## RBSPICEB_Notes
Calibration File Date: |  | 2016-03-06
Calibration Revision |  | 15
Applicable Calibration Products
Product | Description |  | Sector Cadence | Spin Cadence | # Energy Channels
ESRHELT | Electron Rates measured using the RBSPICE Energy Mode: High Energy Resolution Low Time Resolution |  | Multi Sector Group | Multi Spin | 64
ESRLEHT | Electron Rates measured using the RBSPICE Energy Mode: Low Energy Resolution High Time Resolution |  | Single Sector Group | Single Spin | 14
ISRHELT | Ion Rates measured using the RBSPICE Energy Mode: High Energy Resolution Low Time Resolution |  | Multi Sector Group | Multi Spin | 64
TOFxEH | Ion Rates measured using the RBSPICE Species Mode: Hydrogen Only Observations |  | Single Sector Group | Single Spin | 14
TOFxEIon | Ion Rates measured using the RBSPICE Species Mode: All Ion Observations |  | Multi Sector Group | Multi Spin | 64
TOFxEnonH | Ion Rates measured using the RBSPICE Species Mode: Z>1 Observations |  | Multi Sector Group | Single Spin | 20
Note that the TOFxEnonH calibration information is broken into two tabs in this spreadsheet: pre-He extension and post-He extension.
The He extension occurred on: 2015-01-14T23:45:00.0
TOFxPHHHELT | Ion Rates measured using the RBSPICE Species Mode using only the TOF system: Hydrogen Observations High Energy Resolution Low Time Resolution |  | Multi Sector Group | Multi Spin | 32
TOFxPHHLEHT | Ion Rates measured using the RBSPICE Species Mode using only the TOF system: Hydrogen Observations Low Energy Resolution High Time Resolution |  | Single Sector Group | Single Spin | 10
Cadence Explanation: |  | The RBSPICE instrument flight software breaks a single spin into 36 sectors. Accumulation of counts is done winthin a single sector group, Multi sector groups, and Multi Spins
Single Sector Group |  | Accumulation involves one or more sectors but at the currently defined grouping but within the context of a single spin.
Grouping can be identified by examining the file from a Single Sector Group product and seeing the number of sectors used in each measurement
I.e. if the sector number in a file from a single sector group product for a specific spin number has the following values: 0, 3, 6, 9, … then the sector group size is 3
Multi Sector Group |  | Accumulation involves more than one sector group but is still done within the context of a single spin
Single Spin |  | Accumulation is taken over a single spin.
Multi Spin |  | Accumulation is taken over more than one sector group and is done across multiple spins.
Note that a multi Spin measurement still preserves the integrity of the angular measurement since a group of sectors is measured for each of the spins involved in the observation
I.e. if the spin number is 10 and we are at a sector group number of 3 and start our observation in sector 9 then accumulation of counts will occur for sectors 9 through 11 for each of the spins
Field Definitions
SC | Name of spacecraft either RBSPA or RBSPB
ProductType | One of the product types listed in the beginning of the cal file
Telescope | Identifies which of the six telescopes the cal information corresponds
StartUTC | The starting time for this calibration record in UTC string format (CCYY-MM-DDTHH:MM:SS:hhh)
StartET | The starting time for this calibration record in Ephemeris Time using the J2000 epoch
StopUTC | The ending time for this calibration record in UTC string format (CCYY-MM-DDTHH:MM:SS:hhh)
StopET | The ending time for this calibration record in Ephemeris Time using the J2000 epoch
Species | The primary species for which this calibration record is responsive
Note that this does not identify all species that this channel will detect but the species that it is designed to detect.
Some channels are responsive to multiple species and depending upon the situation the primary species differs from time to time.
E.g. TOFxPH products are generally responsive to protons but some of the channels are responsive to Oxygen or Helium although when those species are not present the channel will detect background proton rates
Channel | The energy channel number for this calibration record
E_Low | The bottom energy of the passband in MeV
E_High | The upper energy of the passband in MeV
E_Mid | The calculated midpoint energy of the passband in MeV.
Note that this is not always the geometrical mean since some passbands are more sensitive to lower energies even though they allow for higher energy ranges
G_Small | The small pixel geometrical factor in cm^2*sr. See the RBSPICE Data Handbook for more information about pixel sizes
G_Large | The large pixel geometrical factor in cm^2*sr. See the RBSPICE Data Handbook for more information about pixel sizes
Eff | The efficiency of the energy channel.
Notes | Any specific notes about this energy channel.

## ESR_LEHT
Product Description | Product Name | StartET | StopET | Species | Field Name | Channel | ChName | E_Low (Kev) | E_Hi (Kev) | Mid_E(KeV) | G (cm^2 sr) | G-small | efficiency
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 0 | c00 | 20.635 | 24.7462 | 22.6906 | 0.001675 | 0.000335 | 0.4
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 1 | c01 | 24.7462 | 30.1345 | 27.4404 | 0.001675 | 0.0001675 | 0.6
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 2 | c02 | 30.1345 | 38.0573 | 34.0959 | 0.001675 | 0.00014238 | 0.72
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 3 | c03 | 38.0573 | 47.9656 | 43.0114 | 0.001675 | 0.00012563 | 0.82
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 4 | c04 | 47.9656 | 60.6144 | 54.29 | 0.001675 | 0.00012563 | 0.9
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 5 | c05 | 60.6144 | 73.8454 | 67.2299 | 0.001675 | 0.00012563 | 0.94
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 6 | c06 | 73.8454 | 89.2777 | 81.5615 | 0.001675 | 0.00012563 | 0.97
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 7 | c07 | 89.2777 | 108.0352 | 98.6564 | 0.001675 | 0.00012563 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 8 | c08 | 108.0352 | 138.1469 | 123.0911 | 0.001675 | 0.00012563 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 9 | c09 | 138.1469 | 176.76 | 157.4534 | 0.001675 | 0.00012563 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 10 | c10 | 176.76 | 212.3748 | 194.5674 | 0.001675 | 0.00012563 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 11 | c11 | 212.3748 | 270.4472 | 241.411 | 0.001675 | 0.00012563 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 12 | c12 | 270.4472 | 324.5262 | 297.4867 | 0.001675 | 0.00012563 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 13 | c13 | 324.5262 | 526.9775 | 425.7518 | 0.001675 | 0.00012563 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 0 | c00 | 20.635 | 24.7462 | 22.6906 | 0.0017478 | 0.0017478 | 0.4
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 1 | c01 | 24.7462 | 30.1345 | 27.4404 | 0.0017478 | 0.00013982 | 0.6
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 2 | c02 | 30.1345 | 38.0573 | 34.0959 | 0.0017478 | 9.6127e-05 | 0.72
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 3 | c03 | 38.0573 | 47.9656 | 43.0114 | 0.0017478 | 0.00011361 | 0.82
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 4 | c04 | 47.9656 | 60.6144 | 54.29 | 0.0017478 | 0.00011361 | 0.9
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 5 | c05 | 60.6144 | 73.8454 | 67.2299 | 0.0017478 | 0.00011361 | 0.94
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 6 | c06 | 73.8454 | 89.2777 | 81.5615 | 0.0017478 | 0.00011361 | 0.97
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 7 | c07 | 89.2777 | 108.0352 | 98.6564 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 8 | c08 | 108.0352 | 138.1469 | 123.0911 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 9 | c09 | 138.1469 | 176.76 | 157.4534 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 10 | c10 | 176.76 | 212.3748 | 194.5674 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 11 | c11 | 212.3748 | 270.4472 | 241.411 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 12 | c12 | 270.4472 | 324.5262 | 297.4867 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 13 | c13 | 324.5262 | 526.9775 | 425.7518 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 0 | c00 | 20.635 | 24.7462 | 22.6906 | 0.0014603 | 0.0021905 | 0.4
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 1 | c01 | 24.7462 | 30.1345 | 27.4404 | 0.0014603 | 0.00014603 | 0.6
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 2 | c02 | 30.1345 | 38.0573 | 34.0959 | 0.0014603 | 8.0318e-05 | 0.72
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 3 | c03 | 38.0573 | 47.9656 | 43.0114 | 0.0014603 | 8.7619e-05 | 0.82
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 4 | c04 | 47.9656 | 60.6144 | 54.29 | 0.0014603 | 8.7619e-05 | 0.9
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 5 | c05 | 60.6144 | 73.8454 | 67.2299 | 0.0014603 | 8.7619e-05 | 0.94
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 6 | c06 | 73.8454 | 89.2777 | 81.5615 | 0.0014603 | 8.7619e-05 | 0.97
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 7 | c07 | 89.2777 | 108.0352 | 98.6564 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 8 | c08 | 108.0352 | 138.1469 | 123.0911 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 9 | c09 | 138.1469 | 176.76 | 157.4534 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 10 | c10 | 176.76 | 212.3748 | 194.5674 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 11 | c11 | 212.3748 | 270.4472 | 241.411 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 12 | c12 | 270.4472 | 324.5262 | 297.4867 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 13 | c13 | 324.5262 | 526.9775 | 425.7518 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 0 | c00 | 20.635 | 24.7462 | 22.6906 | 0.0015809 | 0.00023713 | 0.4
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 1 | c01 | 24.7462 | 30.1345 | 27.4404 | 0.0015809 | 0.00013437 | 0.6
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 2 | c02 | 30.1345 | 38.0573 | 34.0959 | 0.0015809 | 0.00011856 | 0.72
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 3 | c03 | 38.0573 | 47.9656 | 43.0114 | 0.0015809 | 9.4851e-05 | 0.82
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 4 | c04 | 47.9656 | 60.6144 | 54.29 | 0.0015809 | 9.4851e-05 | 0.9
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 5 | c05 | 60.6144 | 73.8454 | 67.2299 | 0.0015809 | 9.4851e-05 | 0.94
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 6 | c06 | 73.8454 | 89.2777 | 81.5615 | 0.0015809 | 9.4851e-05 | 0.97
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 7 | c07 | 89.2777 | 108.0352 | 98.6564 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 8 | c08 | 108.0352 | 138.1469 | 123.0911 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 9 | c09 | 138.1469 | 176.76 | 157.4534 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 10 | c10 | 176.76 | 212.3748 | 194.5674 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 11 | c11 | 212.3748 | 270.4472 | 241.411 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 12 | c12 | 270.4472 | 324.5262 | 297.4867 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 13 | c13 | 324.5262 | 526.9775 | 425.7518 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 0 | c00 | 20.635 | 24.7462 | 22.6906 | 0.0016419 | 0.0001724 | 0.4
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 1 | c01 | 24.7462 | 30.1345 | 27.4404 | 0.0016419 | 0.00012314 | 0.6
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 2 | c02 | 30.1345 | 38.0573 | 34.0959 | 0.0016419 | 9.8514e-05 | 0.72
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 3 | c03 | 38.0573 | 47.9656 | 43.0114 | 0.0016419 | 9.8514e-05 | 0.82
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 4 | c04 | 47.9656 | 60.6144 | 54.29 | 0.0016419 | 9.8514e-05 | 0.9
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 5 | c05 | 60.6144 | 73.8454 | 67.2299 | 0.0016419 | 9.8514e-05 | 0.94
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 6 | c06 | 73.8454 | 89.2777 | 81.5615 | 0.0016419 | 9.8514e-05 | 0.97
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 7 | c07 | 89.2777 | 108.0352 | 98.6564 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 8 | c08 | 108.0352 | 138.1469 | 123.0911 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 9 | c09 | 138.1469 | 176.76 | 157.4534 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 10 | c10 | 176.76 | 212.3748 | 194.5674 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 11 | c11 | 212.3748 | 270.4472 | 241.411 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 12 | c12 | 270.4472 | 324.5262 | 297.4867 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 13 | c13 | 324.5262 | 526.9775 | 425.7518 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 0 | c00 | 20.635 | 24.7462 | 22.6906 | 3.1876e-05 | 1.5938e-06 | 0.4 | BLOCKED
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 1 | c01 | 24.7462 | 30.1345 | 27.4404 | 3.1876e-05 | 1.5938e-06 | 0.6 | BLOCKED
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 2 | c02 | 30.1345 | 38.0573 | 34.0959 | 3.1876e-05 | 1.5938e-06 | 0.72 | BLOCKED
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 3 | c03 | 38.0573 | 47.9656 | 43.0114 | 3.1876e-05 | 1.5938e-06 | 0.82 | BLOCKED
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 4 | c04 | 47.9656 | 60.6144 | 54.29 | 3.1876e-05 | 1.5938e-06 | 0.9 | BLOCKED
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 5 | c05 | 60.6144 | 73.8454 | 67.2299 | 3.1876e-05 | 1.5938e-06 | 0.94 | BLOCKED
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 6 | c06 | 73.8454 | 89.2777 | 81.5615 | 3.1876e-05 | 1.5938e-06 | 0.97 | BLOCKED
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 7 | c07 | 89.2777 | 108.0352 | 98.6564 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 8 | c08 | 108.0352 | 138.1469 | 123.0911 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 9 | c09 | 138.1469 | 176.76 | 157.4534 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 10 | c10 | 176.76 | 212.3748 | 194.5674 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 11 | c11 | 212.3748 | 270.4472 | 241.411 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 12 | c12 | 270.4472 | 324.5262 | 297.4867 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - Low Energy Res, High Time Res | ESR_LEHT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 13 | c13 | 324.5262 | 526.9775 | 425.7518 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED

## ESR_HELT
Product Description | Product Name | StartET | StopET | Species | Field Name | Channel | ChName | E_Low (Kev) | E_Hi (Kev) | Mid_E(KeV) | G (cm^2 sr) | G-small | efficiency
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 0 | e00 | 19.7113 | 20.635 | 20.1732 | 0.001675 | 0.0335 | 10.345
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 1 | e01 | 20.635 | 22.0122 | 21.3236 | 0.001675 | 0.001675 | 0.001
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 2 | e02 | 22.0122 | 23.3841 | 22.6982 | 0.001675 | 0.000335 | 0.006
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 3 | e03 | 23.3841 | 24.7462 | 24.0652 | 0.001675 | 0.00019263 | 0.157
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 4 | e04 | 24.7462 | 26.5533 | 25.6498 | 0.001675 | 0.00015075 | 0.452
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 5 | e05 | 26.5533 | 28.3491 | 27.4512 | 0.001675 | 0.0001675 | 0.577
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 6 | e06 | 28.3491 | 30.1345 | 29.2418 | 0.001675 | 0.0001675 | 0.595
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 7 | e07 | 30.1345 | 31.9102 | 31.0224 | 0.001675 | 0.0001675 | 0.639
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 8 | e08 | 31.9102 | 33.6768 | 32.7935 | 0.001675 | 0.00014238 | 0.688
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 9 | e09 | 33.6768 | 35.872 | 34.7744 | 0.001675 | 0.00014238 | 0.881
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 10 | e10 | 35.872 | 38.0573 | 36.9646 | 0.001675 | 0.000134 | 0.588
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 11 | e11 | 38.0573 | 40.2288 | 39.143 | 0.001675 | 0.000134 | 0.97
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 12 | e12 | 40.2288 | 42.8218 | 41.5253 | 0.001675 | 0.00012563 | 0.817
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 13 | e13 | 42.8218 | 45.4004 | 44.1111 | 0.001675 | 0.00012563 | 0.814
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 14 | e14 | 45.4004 | 47.9656 | 46.683 | 0.001675 | 0.00012563 | 0.813
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 15 | e15 | 47.9656 | 50.9413 | 49.4534 | 0.001675 | 0.00012563 | 1.018
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 16 | e16 | 50.9413 | 53.9033 | 52.4223 | 0.001675 | 0.00012563 | 0.8
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 17 | e17 | 53.9033 | 56.8483 | 55.3758 | 0.001675 | 0.00012563 | 1.075
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 18 | e18 | 56.8483 | 60.6144 | 58.7314 | 0.001675 | 0.00012563 | 0.973
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 19 | e19 | 60.6144 | 64.7753 | 62.6949 | 0.001675 | 0.00012563 | 0.881
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 20 | e20 | 64.7753 | 69.3228 | 67.0491 | 0.001675 | 0.00012563 | 1.038
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 21 | e21 | 69.3228 | 73.8454 | 71.5841 | 0.001675 | 0.00012563 | 0.853
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 22 | e22 | 73.8454 | 78.7466 | 76.296 | 0.001675 | 0.00012563 | 1.133
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 23 | e23 | 78.7466 | 84.0285 | 81.3876 | 0.001675 | 0.00012563 | 0.875
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 24 | e24 | 84.0285 | 89.2777 | 86.6531 | 0.001675 | 0.00012563 | 1.018
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 25 | e25 | 89.2777 | 95.2995 | 92.2886 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 26 | e26 | 95.2995 | 101.6858 | 98.4927 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 27 | e27 | 101.6858 | 108.0352 | 104.8605 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 28 | e28 | 108.0352 | 114.743 | 111.3891 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 29 | e29 | 114.743 | 122.1976 | 118.4703 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 30 | e30 | 122.1976 | 130.0008 | 126.0992 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 31 | e31 | 130.0008 | 138.1469 | 134.0739 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 32 | e32 | 138.1469 | 147.0186 | 142.5827 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 33 | e33 | 147.0186 | 156.2248 | 151.6217 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 34 | e34 | 156.2248 | 166.1401 | 161.1825 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 35 | e35 | 166.1401 | 176.76 | 171.4501 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 36 | e36 | 176.76 | 187.6964 | 182.2282 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 37 | e37 | 187.6964 | 199.6957 | 193.696 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 38 | e38 | 199.6957 | 212.3748 | 206.0353 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 39 | e39 | 212.3748 | 225.3535 | 218.8642 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 40 | e40 | 225.3535 | 239.3662 | 232.3599 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 41 | e41 | 239.3662 | 254.4023 | 246.8843 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 42 | e42 | 254.4023 | 270.4472 | 262.4247 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 43 | e43 | 270.4472 | 287.491 | 278.9691 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 44 | e44 | 287.491 | 305.5203 | 296.5057 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 45 | e45 | 305.5203 | 324.5262 | 315.0232 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 46 | e46 | 324.5262 | 344.8509 | 334.6886 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 47 | e47 | 344.8509 | 366.4828 | 355.6669 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 48 | e48 | 366.4828 | 389.4026 | 377.9427 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 49 | e49 | 389.4026 | 413.9488 | 401.6757 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 50 | e50 | 413.9488 | 439.7528 | 426.8508 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 51 | e51 | 439.7528 | 467.1467 | 453.4498 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 52 | e52 | 467.1467 | 496.1143 | 481.6305 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 53 | e53 | 496.1143 | 526.9775 | 511.5459 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 54 | e54 | 526.9775 | 560.0573 | 543.5174 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 55 | e55 | 560.0573 | 594.9878 | 577.5225 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 56 | e56 | 594.9878 | 632.0842 | 613.536 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 57 | e57 | 632.0842 | 671.3239 | 651.704 | 0.001675 | 0.00012563 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 58 | e58 | 671.3239 | 712.6814 | 692.0026 | 0.001675 | 0.00012563 | 0.9
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 59 | e59 | 712.6814 | 756.4661 | 734.5737 | 0.001675 | 0.00012563 | 0.8
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 60 | e60 | 756.4661 | 802.6511 | 779.5586 | 0.001675 | 0.00012563 | 0.6
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 61 | e61 | 802.6511 | 851.5375 | 827.0943 | 0.001675 | 0.00012563 | 0.4
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 62 | e62 | 851.5375 | 903.0973 | 877.3174 | 0.001675 | 0.00012563 | 0.2
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T0_Counts | 63 | e63 | 903.0973 | 974.3854 | 938.7414 | 0.001675 | 0.00012563 | 0.05
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 0 | e00 | 19.7113 | 20.635 | 20.1732 | 0.0017478 | 0.034955 | 9.525
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 1 | e01 | 20.635 | 22.0122 | 21.3236 | 0.0017478 | 0.034955 | 0.004
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 2 | e02 | 22.0122 | 23.3841 | 22.6982 | 0.0017478 | 0.0026217 | 0.019
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 3 | e03 | 23.3841 | 24.7462 | 24.0652 | 0.0017478 | 0.00034955 | 0.273
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 4 | e04 | 24.7462 | 26.5533 | 25.6498 | 0.0017478 | 0.00017478 | 0.513
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 5 | e05 | 26.5533 | 28.3491 | 27.4512 | 0.0017478 | 0.00013982 | 0.496
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 6 | e06 | 28.3491 | 30.1345 | 29.2418 | 0.0017478 | 0.00013108 | 0.643
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 7 | e07 | 30.1345 | 31.9102 | 31.0224 | 0.0017478 | 0.00013108 | 0.617
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 8 | e08 | 31.9102 | 33.6768 | 32.7935 | 0.0017478 | 0.00012234 | 0.733
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 9 | e09 | 33.6768 | 35.872 | 34.7744 | 0.0017478 | 9.6127e-05 | 0.869
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 10 | e10 | 35.872 | 38.0573 | 36.9646 | 0.0017478 | 9.6127e-05 | 0.592
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 11 | e11 | 38.0573 | 40.2288 | 39.143 | 0.0017478 | 9.6127e-05 | 0.994
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 12 | e12 | 40.2288 | 42.8218 | 41.5253 | 0.0017478 | 9.6127e-05 | 0.796
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 13 | e13 | 42.8218 | 45.4004 | 44.1111 | 0.0017478 | 0.00011361 | 0.807
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 14 | e14 | 45.4004 | 47.9656 | 46.683 | 0.0017478 | 0.00011361 | 0.862
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 15 | e15 | 47.9656 | 50.9413 | 49.4534 | 0.0017478 | 0.00011361 | 1.008
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 16 | e16 | 50.9413 | 53.9033 | 52.4223 | 0.0017478 | 0.00011361 | 0.811
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 17 | e17 | 53.9033 | 56.8483 | 55.3758 | 0.0017478 | 0.00011361 | 1.075
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 18 | e18 | 56.8483 | 60.6144 | 58.7314 | 0.0017478 | 0.00011361 | 1.032
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 19 | e19 | 60.6144 | 64.7753 | 62.6949 | 0.0017478 | 0.00011361 | 0.845
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 20 | e20 | 64.7753 | 69.3228 | 67.0491 | 0.0017478 | 0.00011361 | 1.077
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 21 | e21 | 69.3228 | 73.8454 | 71.5841 | 0.0017478 | 0.00011361 | 0.875
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 22 | e22 | 73.8454 | 78.7466 | 76.296 | 0.0017478 | 0.00011361 | 1.161
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 23 | e23 | 78.7466 | 84.0285 | 81.3876 | 0.0017478 | 0.00011361 | 0.862
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 24 | e24 | 84.0285 | 89.2777 | 86.6531 | 0.0017478 | 0.00011361 | 1.046
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 25 | e25 | 89.2777 | 95.2995 | 92.2886 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 26 | e26 | 95.2995 | 101.6858 | 98.4927 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 27 | e27 | 101.6858 | 108.0352 | 104.8605 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 28 | e28 | 108.0352 | 114.743 | 111.3891 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 29 | e29 | 114.743 | 122.1976 | 118.4703 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 30 | e30 | 122.1976 | 130.0008 | 126.0992 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 31 | e31 | 130.0008 | 138.1469 | 134.0739 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 32 | e32 | 138.1469 | 147.0186 | 142.5827 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 33 | e33 | 147.0186 | 156.2248 | 151.6217 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 34 | e34 | 156.2248 | 166.1401 | 161.1825 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 35 | e35 | 166.1401 | 176.76 | 171.4501 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 36 | e36 | 176.76 | 187.6964 | 182.2282 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 37 | e37 | 187.6964 | 199.6957 | 193.696 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 38 | e38 | 199.6957 | 212.3748 | 206.0353 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 39 | e39 | 212.3748 | 225.3535 | 218.8642 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 40 | e40 | 225.3535 | 239.3662 | 232.3599 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 41 | e41 | 239.3662 | 254.4023 | 246.8843 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 42 | e42 | 254.4023 | 270.4472 | 262.4247 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 43 | e43 | 270.4472 | 287.491 | 278.9691 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 44 | e44 | 287.491 | 305.5203 | 296.5057 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 45 | e45 | 305.5203 | 324.5262 | 315.0232 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 46 | e46 | 324.5262 | 344.8509 | 334.6886 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 47 | e47 | 344.8509 | 366.4828 | 355.6669 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 48 | e48 | 366.4828 | 389.4026 | 377.9427 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 49 | e49 | 389.4026 | 413.9488 | 401.6757 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 50 | e50 | 413.9488 | 439.7528 | 426.8508 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 51 | e51 | 439.7528 | 467.1467 | 453.4498 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 52 | e52 | 467.1467 | 496.1143 | 481.6305 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 53 | e53 | 496.1143 | 526.9775 | 511.5459 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 54 | e54 | 526.9775 | 560.0573 | 543.5174 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 55 | e55 | 560.0573 | 594.9878 | 577.5225 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 56 | e56 | 594.9878 | 632.0842 | 613.536 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 57 | e57 | 632.0842 | 671.3239 | 651.704 | 0.0017478 | 0.00011361 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 58 | e58 | 671.3239 | 712.6814 | 692.0026 | 0.0017478 | 0.00011361 | 0.9
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 59 | e59 | 712.6814 | 756.4661 | 734.5737 | 0.0017478 | 0.00011361 | 0.8
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 60 | e60 | 756.4661 | 802.6511 | 779.5586 | 0.0017478 | 0.00011361 | 0.6
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 61 | e61 | 802.6511 | 851.5375 | 827.0943 | 0.0017478 | 0.00011361 | 0.4
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 62 | e62 | 851.5375 | 903.0973 | 877.3174 | 0.0017478 | 0.00011361 | 0.2
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T1_Counts | 63 | e63 | 903.0973 | 974.3854 | 938.7414 | 0.0017478 | 0.00011361 | 0.05
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 0 | e00 | 19.7113 | 20.635 | 20.1732 | 0.0014603 | 0.029206 | 8.408
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 1 | e01 | 20.635 | 22.0122 | 21.3236 | 0.0014603 | 0.058413 | 0.001
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 2 | e02 | 22.0122 | 23.3841 | 22.6982 | 0.0014603 | 0.0036508 | 0.004
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 3 | e03 | 23.3841 | 24.7462 | 24.0652 | 0.0014603 | 0.0004381 | 0.143
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 4 | e04 | 24.7462 | 26.5533 | 25.6498 | 0.0014603 | 0.00017524 | 0.284
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 5 | e05 | 26.5533 | 28.3491 | 27.4512 | 0.0014603 | 0.00014603 | 0.467
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 6 | e06 | 28.3491 | 30.1345 | 29.2418 | 0.0014603 | 0.00010952 | 0.491
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 7 | e07 | 30.1345 | 31.9102 | 31.0224 | 0.0014603 | 9.4921e-05 | 0.7
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 8 | e08 | 31.9102 | 33.6768 | 32.7935 | 0.0014603 | 8.0318e-05 | 0.702
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 9 | e09 | 33.6768 | 35.872 | 34.7744 | 0.0014603 | 8.0318e-05 | 0.992
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 10 | e10 | 35.872 | 38.0573 | 36.9646 | 0.0014603 | 8.7619e-05 | 0.704
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 11 | e11 | 38.0573 | 40.2288 | 39.143 | 0.0014603 | 8.7619e-05 | 1.006
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 12 | e12 | 40.2288 | 42.8218 | 41.5253 | 0.0014603 | 9.4921e-05 | 0.85
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 13 | e13 | 42.8218 | 45.4004 | 44.1111 | 0.0014603 | 8.7619e-05 | 0.721
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 14 | e14 | 45.4004 | 47.9656 | 46.683 | 0.0014603 | 8.7619e-05 | 0.839
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 15 | e15 | 47.9656 | 50.9413 | 49.4534 | 0.0014603 | 8.7619e-05 | 1.059
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 16 | e16 | 50.9413 | 53.9033 | 52.4223 | 0.0014603 | 8.7619e-05 | 0.721
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 17 | e17 | 53.9033 | 56.8483 | 55.3758 | 0.0014603 | 8.7619e-05 | 1.04
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 18 | e18 | 56.8483 | 60.6144 | 58.7314 | 0.0014603 | 8.7619e-05 | 1.077
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 19 | e19 | 60.6144 | 64.7753 | 62.6949 | 0.0014603 | 8.7619e-05 | 0.942
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 20 | e20 | 64.7753 | 69.3228 | 67.0491 | 0.0014603 | 8.7619e-05 | 1.026
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 21 | e21 | 69.3228 | 73.8454 | 71.5841 | 0.0014603 | 8.7619e-05 | 0.888
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 22 | e22 | 73.8454 | 78.7466 | 76.296 | 0.0014603 | 8.7619e-05 | 1.107
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 23 | e23 | 78.7466 | 84.0285 | 81.3876 | 0.0014603 | 8.7619e-05 | 0.888
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 24 | e24 | 84.0285 | 89.2777 | 86.6531 | 0.0014603 | 8.7619e-05 | 1.018
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 25 | e25 | 89.2777 | 95.2995 | 92.2886 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 26 | e26 | 95.2995 | 101.6858 | 98.4927 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 27 | e27 | 101.6858 | 108.0352 | 104.8605 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 28 | e28 | 108.0352 | 114.743 | 111.3891 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 29 | e29 | 114.743 | 122.1976 | 118.4703 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 30 | e30 | 122.1976 | 130.0008 | 126.0992 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 31 | e31 | 130.0008 | 138.1469 | 134.0739 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 32 | e32 | 138.1469 | 147.0186 | 142.5827 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 33 | e33 | 147.0186 | 156.2248 | 151.6217 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 34 | e34 | 156.2248 | 166.1401 | 161.1825 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 35 | e35 | 166.1401 | 176.76 | 171.4501 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 36 | e36 | 176.76 | 187.6964 | 182.2282 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 37 | e37 | 187.6964 | 199.6957 | 193.696 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 38 | e38 | 199.6957 | 212.3748 | 206.0353 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 39 | e39 | 212.3748 | 225.3535 | 218.8642 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 40 | e40 | 225.3535 | 239.3662 | 232.3599 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 41 | e41 | 239.3662 | 254.4023 | 246.8843 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 42 | e42 | 254.4023 | 270.4472 | 262.4247 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 43 | e43 | 270.4472 | 287.491 | 278.9691 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 44 | e44 | 287.491 | 305.5203 | 296.5057 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 45 | e45 | 305.5203 | 324.5262 | 315.0232 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 46 | e46 | 324.5262 | 344.8509 | 334.6886 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 47 | e47 | 344.8509 | 366.4828 | 355.6669 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 48 | e48 | 366.4828 | 389.4026 | 377.9427 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 49 | e49 | 389.4026 | 413.9488 | 401.6757 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 50 | e50 | 413.9488 | 439.7528 | 426.8508 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 51 | e51 | 439.7528 | 467.1467 | 453.4498 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 52 | e52 | 467.1467 | 496.1143 | 481.6305 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 53 | e53 | 496.1143 | 526.9775 | 511.5459 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 54 | e54 | 526.9775 | 560.0573 | 543.5174 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 55 | e55 | 560.0573 | 594.9878 | 577.5225 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 56 | e56 | 594.9878 | 632.0842 | 613.536 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 57 | e57 | 632.0842 | 671.3239 | 651.704 | 0.0014603 | 8.7619e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 58 | e58 | 671.3239 | 712.6814 | 692.0026 | 0.0014603 | 8.7619e-05 | 0.9
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 59 | e59 | 712.6814 | 756.4661 | 734.5737 | 0.0014603 | 8.7619e-05 | 0.8
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 60 | e60 | 756.4661 | 802.6511 | 779.5586 | 0.0014603 | 8.7619e-05 | 0.6
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 61 | e61 | 802.6511 | 851.5375 | 827.0943 | 0.0014603 | 8.7619e-05 | 0.4
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 62 | e62 | 851.5375 | 903.0973 | 877.3174 | 0.0014603 | 8.7619e-05 | 0.2
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T2_Counts | 63 | e63 | 903.0973 | 974.3854 | 938.7414 | 0.0014603 | 8.7619e-05 | 0.05
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 0 | e00 | 19.7113 | 20.635 | 20.1732 | 0.0015809 | 0.031617 | 8.981
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 1 | e01 | 20.635 | 22.0122 | 21.3236 | 0.0015809 | 0.0011856 | 0.006
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 2 | e02 | 22.0122 | 23.3841 | 22.6982 | 0.0015809 | 0.00030827 | 0.035
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 3 | e03 | 23.3841 | 24.7462 | 24.0652 | 0.0015809 | 0.00016599 | 0.37
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 4 | e04 | 24.7462 | 26.5533 | 25.6498 | 0.0015809 | 0.00014228 | 0.554
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 5 | e05 | 26.5533 | 28.3491 | 27.4512 | 0.0015809 | 0.00013437 | 0.5
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 6 | e06 | 28.3491 | 30.1345 | 29.2418 | 0.0015809 | 0.00013437 | 0.65
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 7 | e07 | 30.1345 | 31.9102 | 31.0224 | 0.0015809 | 0.00013437 | 0.62
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 8 | e08 | 31.9102 | 33.6768 | 32.7935 | 0.0015809 | 0.00012647 | 0.744
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 9 | e09 | 33.6768 | 35.872 | 34.7744 | 0.0015809 | 0.00011856 | 0.872
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 10 | e10 | 35.872 | 38.0573 | 36.9646 | 0.0015809 | 0.00011856 | 0.593
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 11 | e11 | 38.0573 | 40.2288 | 39.143 | 0.0015809 | 0.00011856 | 0.987
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 12 | e12 | 40.2288 | 42.8218 | 41.5253 | 0.0015809 | 9.4851e-05 | 0.812
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 13 | e13 | 42.8218 | 45.4004 | 44.1111 | 0.0015809 | 9.4851e-05 | 0.803
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 14 | e14 | 45.4004 | 47.9656 | 46.683 | 0.0015809 | 9.4851e-05 | 0.834
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 15 | e15 | 47.9656 | 50.9413 | 49.4534 | 0.0015809 | 9.4851e-05 | 0.863
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 16 | e16 | 50.9413 | 53.9033 | 52.4223 | 0.0015809 | 9.4851e-05 | 0.826
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 17 | e17 | 53.9033 | 56.8483 | 55.3758 | 0.0015809 | 9.4851e-05 | 1.031
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 18 | e18 | 56.8483 | 60.6144 | 58.7314 | 0.0015809 | 9.4851e-05 | 1.047
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 19 | e19 | 60.6144 | 64.7753 | 62.6949 | 0.0015809 | 9.4851e-05 | 0.966
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 20 | e20 | 64.7753 | 69.3228 | 67.0491 | 0.0015809 | 9.4851e-05 | 1.069
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 21 | e21 | 69.3228 | 73.8454 | 71.5841 | 0.0015809 | 9.4851e-05 | 0.891
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 22 | e22 | 73.8454 | 78.7466 | 76.296 | 0.0015809 | 9.4851e-05 | 1.165
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 23 | e23 | 78.7466 | 84.0285 | 81.3876 | 0.0015809 | 9.4851e-05 | 0.877
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 24 | e24 | 84.0285 | 89.2777 | 86.6531 | 0.0015809 | 9.4851e-05 | 1.059
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 25 | e25 | 89.2777 | 95.2995 | 92.2886 | 0.0015809 | 9.4851e-05 | 1.021
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 26 | e26 | 95.2995 | 101.6858 | 98.4927 | 0.0015809 | 9.4851e-05 | 0.933
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 27 | e27 | 101.6858 | 108.0352 | 104.8605 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 28 | e28 | 108.0352 | 114.743 | 111.3891 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 29 | e29 | 114.743 | 122.1976 | 118.4703 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 30 | e30 | 122.1976 | 130.0008 | 126.0992 | 0.0015809 | 9.4851e-05 | 0.909
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 31 | e31 | 130.0008 | 138.1469 | 134.0739 | 0.0015809 | 9.4851e-05 | 0.926
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 32 | e32 | 138.1469 | 147.0186 | 142.5827 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 33 | e33 | 147.0186 | 156.2248 | 151.6217 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 34 | e34 | 156.2248 | 166.1401 | 161.1825 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 35 | e35 | 166.1401 | 176.76 | 171.4501 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 36 | e36 | 176.76 | 187.6964 | 182.2282 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 37 | e37 | 187.6964 | 199.6957 | 193.696 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 38 | e38 | 199.6957 | 212.3748 | 206.0353 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 39 | e39 | 212.3748 | 225.3535 | 218.8642 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 40 | e40 | 225.3535 | 239.3662 | 232.3599 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 41 | e41 | 239.3662 | 254.4023 | 246.8843 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 42 | e42 | 254.4023 | 270.4472 | 262.4247 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 43 | e43 | 270.4472 | 287.491 | 278.9691 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 44 | e44 | 287.491 | 305.5203 | 296.5057 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 45 | e45 | 305.5203 | 324.5262 | 315.0232 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 46 | e46 | 324.5262 | 344.8509 | 334.6886 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 47 | e47 | 344.8509 | 366.4828 | 355.6669 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 48 | e48 | 366.4828 | 389.4026 | 377.9427 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 49 | e49 | 389.4026 | 413.9488 | 401.6757 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 50 | e50 | 413.9488 | 439.7528 | 426.8508 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 51 | e51 | 439.7528 | 467.1467 | 453.4498 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 52 | e52 | 467.1467 | 496.1143 | 481.6305 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 53 | e53 | 496.1143 | 526.9775 | 511.5459 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 54 | e54 | 526.9775 | 560.0573 | 543.5174 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 55 | e55 | 560.0573 | 594.9878 | 577.5225 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 56 | e56 | 594.9878 | 632.0842 | 613.536 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 57 | e57 | 632.0842 | 671.3239 | 651.704 | 0.0015809 | 9.4851e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 58 | e58 | 671.3239 | 712.6814 | 692.0026 | 0.0015809 | 9.4851e-05 | 0.9
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 59 | e59 | 712.6814 | 756.4661 | 734.5737 | 0.0015809 | 9.4851e-05 | 0.8
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 60 | e60 | 756.4661 | 802.6511 | 779.5586 | 0.0015809 | 9.4851e-05 | 0.6
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 61 | e61 | 802.6511 | 851.5375 | 827.0943 | 0.0015809 | 9.4851e-05 | 0.4
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 62 | e62 | 851.5375 | 903.0973 | 877.3174 | 0.0015809 | 9.4851e-05 | 0.2
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T3_Counts | 63 | e63 | 903.0973 | 974.3854 | 938.7414 | 0.0015809 | 9.4851e-05 | 0.05
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 0 | e00 | 19.7113 | 20.635 | 20.1732 | 0.0016419 | 0.032838 | 9.83
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 1 | e01 | 20.635 | 22.0122 | 21.3236 | 0.0016419 | 0.00049257 | 0.5
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 2 | e02 | 22.0122 | 23.3841 | 22.6982 | 0.0016419 | 0.0001724 | 0.335
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 3 | e03 | 23.3841 | 24.7462 | 24.0652 | 0.0016419 | 0.00014777 | 0.663
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 4 | e04 | 24.7462 | 26.5533 | 25.6498 | 0.0016419 | 0.00013135 | 0.556
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 5 | e05 | 26.5533 | 28.3491 | 27.4512 | 0.0016419 | 0.00012314 | 0.56
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 6 | e06 | 28.3491 | 30.1345 | 29.2418 | 0.0016419 | 0.00011493 | 0.621
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 7 | e07 | 30.1345 | 31.9102 | 31.0224 | 0.0016419 | 0.00010672 | 0.675
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 8 | e08 | 31.9102 | 33.6768 | 32.7935 | 0.0016419 | 9.8514e-05 | 0.723
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 9 | e09 | 33.6768 | 35.872 | 34.7744 | 0.0016419 | 9.8514e-05 | 0.855
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 10 | e10 | 35.872 | 38.0573 | 36.9646 | 0.0016419 | 9.8514e-05 | 0.638
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 11 | e11 | 38.0573 | 40.2288 | 39.143 | 0.0016419 | 9.8514e-05 | 0.783
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 12 | e12 | 40.2288 | 42.8218 | 41.5253 | 0.0016419 | 9.8514e-05 | 0.684
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 13 | e13 | 42.8218 | 45.4004 | 44.1111 | 0.0016419 | 9.8514e-05 | 0.881
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 14 | e14 | 45.4004 | 47.9656 | 46.683 | 0.0016419 | 9.8514e-05 | 0.801
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 15 | e15 | 47.9656 | 50.9413 | 49.4534 | 0.0016419 | 9.8514e-05 | 0.998
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 16 | e16 | 50.9413 | 53.9033 | 52.4223 | 0.0016419 | 9.8514e-05 | 0.764
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 17 | e17 | 53.9033 | 56.8483 | 55.3758 | 0.0016419 | 9.8514e-05 | 1.051
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 18 | e18 | 56.8483 | 60.6144 | 58.7314 | 0.0016419 | 9.8514e-05 | 1.059
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 19 | e19 | 60.6144 | 64.7753 | 62.6949 | 0.0016419 | 9.8514e-05 | 0.967
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 20 | e20 | 64.7753 | 69.3228 | 67.0491 | 0.0016419 | 9.8514e-05 | 1.066
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 21 | e21 | 69.3228 | 73.8454 | 71.5841 | 0.0016419 | 9.8514e-05 | 0.905
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 22 | e22 | 73.8454 | 78.7466 | 76.296 | 0.0016419 | 9.8514e-05 | 1.16
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 23 | e23 | 78.7466 | 84.0285 | 81.3876 | 0.0016419 | 9.8514e-05 | 0.9
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 24 | e24 | 84.0285 | 89.2777 | 86.6531 | 0.0016419 | 9.8514e-05 | 1.078
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 25 | e25 | 89.2777 | 95.2995 | 92.2886 | 0.0016419 | 9.8514e-05 | 1.042
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 26 | e26 | 95.2995 | 101.6858 | 98.4927 | 0.0016419 | 9.8514e-05 | 0.957
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 27 | e27 | 101.6858 | 108.0352 | 104.8605 | 0.0016419 | 9.8514e-05 | 0.939
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 28 | e28 | 108.0352 | 114.743 | 111.3891 | 0.0016419 | 9.8514e-05 | 0.967
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 29 | e29 | 114.743 | 122.1976 | 118.4703 | 0.0016419 | 9.8514e-05 | 0.95
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 30 | e30 | 122.1976 | 130.0008 | 126.0992 | 0.0016419 | 9.8514e-05 | 0.839
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 31 | e31 | 130.0008 | 138.1469 | 134.0739 | 0.0016419 | 9.8514e-05 | 0.944
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 32 | e32 | 138.1469 | 147.0186 | 142.5827 | 0.0016419 | 9.8514e-05 | 0.929
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 33 | e33 | 147.0186 | 156.2248 | 151.6217 | 0.0016419 | 9.8514e-05 | 0.915
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 34 | e34 | 156.2248 | 166.1401 | 161.1825 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 35 | e35 | 166.1401 | 176.76 | 171.4501 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 36 | e36 | 176.76 | 187.6964 | 182.2282 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 37 | e37 | 187.6964 | 199.6957 | 193.696 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 38 | e38 | 199.6957 | 212.3748 | 206.0353 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 39 | e39 | 212.3748 | 225.3535 | 218.8642 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 40 | e40 | 225.3535 | 239.3662 | 232.3599 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 41 | e41 | 239.3662 | 254.4023 | 246.8843 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 42 | e42 | 254.4023 | 270.4472 | 262.4247 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 43 | e43 | 270.4472 | 287.491 | 278.9691 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 44 | e44 | 287.491 | 305.5203 | 296.5057 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 45 | e45 | 305.5203 | 324.5262 | 315.0232 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 46 | e46 | 324.5262 | 344.8509 | 334.6886 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 47 | e47 | 344.8509 | 366.4828 | 355.6669 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 48 | e48 | 366.4828 | 389.4026 | 377.9427 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 49 | e49 | 389.4026 | 413.9488 | 401.6757 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 50 | e50 | 413.9488 | 439.7528 | 426.8508 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 51 | e51 | 439.7528 | 467.1467 | 453.4498 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 52 | e52 | 467.1467 | 496.1143 | 481.6305 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 53 | e53 | 496.1143 | 526.9775 | 511.5459 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 54 | e54 | 526.9775 | 560.0573 | 543.5174 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 55 | e55 | 560.0573 | 594.9878 | 577.5225 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 56 | e56 | 594.9878 | 632.0842 | 613.536 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 57 | e57 | 632.0842 | 671.3239 | 651.704 | 0.0016419 | 9.8514e-05 | 1
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 58 | e58 | 671.3239 | 712.6814 | 692.0026 | 0.0016419 | 9.8514e-05 | 0.9
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 59 | e59 | 712.6814 | 756.4661 | 734.5737 | 0.0016419 | 9.8514e-05 | 0.8
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 60 | e60 | 756.4661 | 802.6511 | 779.5586 | 0.0016419 | 9.8514e-05 | 0.6
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 61 | e61 | 802.6511 | 851.5375 | 827.0943 | 0.0016419 | 9.8514e-05 | 0.4
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 62 | e62 | 851.5375 | 903.0973 | 877.3174 | 0.0016419 | 9.8514e-05 | 0.2
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T4_Counts | 63 | e63 | 903.0973 | 974.3854 | 938.7414 | 0.0016419 | 9.8514e-05 | 0.05
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 0 | e00 | 19.7113 | 20.635 | 20.1732 | 3.1876e-05 | 1.5938e-06 | 10.345 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 1 | e01 | 20.635 | 22.0122 | 21.3236 | 3.1876e-05 | 1.5938e-06 | 0.001 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 2 | e02 | 22.0122 | 23.3841 | 22.6982 | 3.1876e-05 | 1.5938e-06 | 0.006 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 3 | e03 | 23.3841 | 24.7462 | 24.0652 | 3.1876e-05 | 1.5938e-06 | 0.157 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 4 | e04 | 24.7462 | 26.5533 | 25.6498 | 3.1876e-05 | 1.5938e-06 | 0.452 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 5 | e05 | 26.5533 | 28.3491 | 27.4512 | 3.1876e-05 | 1.5938e-06 | 0.577 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 6 | e06 | 28.3491 | 30.1345 | 29.2418 | 3.1876e-05 | 1.5938e-06 | 0.595 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 7 | e07 | 30.1345 | 31.9102 | 31.0224 | 3.1876e-05 | 1.5938e-06 | 0.639 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 8 | e08 | 31.9102 | 33.6768 | 32.7935 | 3.1876e-05 | 1.5938e-06 | 0.688 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 9 | e09 | 33.6768 | 35.872 | 34.7744 | 3.1876e-05 | 1.5938e-06 | 0.881 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 10 | e10 | 35.872 | 38.0573 | 36.9646 | 3.1876e-05 | 1.5938e-06 | 0.588 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 11 | e11 | 38.0573 | 40.2288 | 39.143 | 3.1876e-05 | 1.5938e-06 | 0.97 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 12 | e12 | 40.2288 | 42.8218 | 41.5253 | 3.1876e-05 | 1.5938e-06 | 0.817 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 13 | e13 | 42.8218 | 45.4004 | 44.1111 | 3.1876e-05 | 1.5938e-06 | 0.814 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 14 | e14 | 45.4004 | 47.9656 | 46.683 | 3.1876e-05 | 1.5938e-06 | 0.813 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 15 | e15 | 47.9656 | 50.9413 | 49.4534 | 3.1876e-05 | 1.5938e-06 | 1.018 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 16 | e16 | 50.9413 | 53.9033 | 52.4223 | 3.1876e-05 | 1.5938e-06 | 0.8 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 17 | e17 | 53.9033 | 56.8483 | 55.3758 | 3.1876e-05 | 1.5938e-06 | 1.075 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 18 | e18 | 56.8483 | 60.6144 | 58.7314 | 3.1876e-05 | 1.5938e-06 | 0.973 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 19 | e19 | 60.6144 | 64.7753 | 62.6949 | 3.1876e-05 | 1.5938e-06 | 0.881 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 20 | e20 | 64.7753 | 69.3228 | 67.0491 | 3.1876e-05 | 1.5938e-06 | 1.038 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 21 | e21 | 69.3228 | 73.8454 | 71.5841 | 3.1876e-05 | 1.5938e-06 | 0.853 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 22 | e22 | 73.8454 | 78.7466 | 76.296 | 3.1876e-05 | 1.5938e-06 | 1.133 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 23 | e23 | 78.7466 | 84.0285 | 81.3876 | 3.1876e-05 | 1.5938e-06 | 0.875 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 24 | e24 | 84.0285 | 89.2777 | 86.6531 | 3.1876e-05 | 1.5938e-06 | 1.018 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 25 | e25 | 89.2777 | 95.2995 | 92.2886 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 26 | e26 | 95.2995 | 101.6858 | 98.4927 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 27 | e27 | 101.6858 | 108.0352 | 104.8605 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 28 | e28 | 108.0352 | 114.743 | 111.3891 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 29 | e29 | 114.743 | 122.1976 | 118.4703 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 30 | e30 | 122.1976 | 130.0008 | 126.0992 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 31 | e31 | 130.0008 | 138.1469 | 134.0739 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 32 | e32 | 138.1469 | 147.0186 | 142.5827 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 33 | e33 | 147.0186 | 156.2248 | 151.6217 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 34 | e34 | 156.2248 | 166.1401 | 161.1825 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 35 | e35 | 166.1401 | 176.76 | 171.4501 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 36 | e36 | 176.76 | 187.6964 | 182.2282 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 37 | e37 | 187.6964 | 199.6957 | 193.696 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 38 | e38 | 199.6957 | 212.3748 | 206.0353 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 39 | e39 | 212.3748 | 225.3535 | 218.8642 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 40 | e40 | 225.3535 | 239.3662 | 232.3599 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 41 | e41 | 239.3662 | 254.4023 | 246.8843 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 42 | e42 | 254.4023 | 270.4472 | 262.4247 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 43 | e43 | 270.4472 | 287.491 | 278.9691 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 44 | e44 | 287.491 | 305.5203 | 296.5057 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 45 | e45 | 305.5203 | 324.5262 | 315.0232 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 46 | e46 | 324.5262 | 344.8509 | 334.6886 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 47 | e47 | 344.8509 | 366.4828 | 355.6669 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 48 | e48 | 366.4828 | 389.4026 | 377.9427 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 49 | e49 | 389.4026 | 413.9488 | 401.6757 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 50 | e50 | 413.9488 | 439.7528 | 426.8508 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 51 | e51 | 439.7528 | 467.1467 | 453.4498 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 52 | e52 | 467.1467 | 496.1143 | 481.6305 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 53 | e53 | 496.1143 | 526.9775 | 511.5459 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 54 | e54 | 526.9775 | 560.0573 | 543.5174 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 55 | e55 | 560.0573 | 594.9878 | 577.5225 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 56 | e56 | 594.9878 | 632.0842 | 613.536 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 57 | e57 | 632.0842 | 671.3239 | 651.704 | 3.1876e-05 | 1.5938e-06 | 1 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 58 | e58 | 671.3239 | 712.6814 | 692.0026 | 3.1876e-05 | 1.5938e-06 | 0.9 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 59 | e59 | 712.6814 | 756.4661 | 734.5737 | 3.1876e-05 | 1.5938e-06 | 0.8 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 60 | e60 | 756.4661 | 802.6511 | 779.5586 | 3.1876e-05 | 1.5938e-06 | 0.6 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 61 | e61 | 802.6511 | 851.5375 | 827.0943 | 3.1876e-05 | 1.5938e-06 | 0.4 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 62 | e62 | 851.5375 | 903.0973 | 877.3174 | 3.1876e-05 | 1.5938e-06 | 0.2 | BLOCKED
Electron Species Rate - High Energy Res, Low Time Res | ESR_HELT | 347112066.1839 | 662731267.1829 | e | T5_Counts | 63 | e63 | 903.0973 | 974.3854 | 938.7414 | 3.1876e-05 | 1.5938e-06 | 0.05 | BLOCKED

## ISR_HELT
Product Description | Product Name | StartET | StopET | Species | Field Name | Channel | ChName | E_Low (Kev) | E_Hi (Kev) | Mid_E(KeV) | G (cm^2 sr) | G-small | efficiency
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 0 | i00 | 39.2625 | 43.9398 | 41.6011 | 0.0015011 | 0.00046244 | 0.00020833
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 1 | i01 | 43.9398 | 48.4219 | 46.1809 | 0.0015011 | 0.00078955 | 0.06
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 2 | i02 | 48.4219 | 52.7317 | 50.5768 | 0.0015011 | 0.00014298 | 0.3
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 3 | i03 | 52.7317 | 57.9075 | 55.3196 | 0.0015011 | 0.00013154 | 0.4
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 4 | i04 | 57.9075 | 64.8216 | 61.3645 | 0.0015011 | 0.00011268 | 0.43333
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 5 | i05 | 64.8216 | 70.4957 | 67.6587 | 0.0015011 | 0.00011433 | 0.7
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 6 | i06 | 70.4957 | 77.7766 | 74.1361 | 0.0015011 | 0.00010324 | 0.75
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 7 | i07 | 77.7766 | 85.6695 | 81.723 | 0.0015011 | 0.0001015 | 0.8
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 8 | i08 | 85.6695 | 94.1679 | 89.9187 | 0.0015011 | 0.00011014 | 0.82
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 9 | i09 | 94.1679 | 104.0937 | 99.1308 | 0.0015011 | 9.9803e-05 | 0.85
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 10 | i10 | 104.0937 | 114.5947 | 109.3442 | 0.0015011 | 9.4565e-05 | 0.9
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 11 | i11 | 114.5947 | 126.4362 | 120.5154 | 0.0015011 | 9.4905e-05 | 0.93
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 12 | i12 | 126.4362 | 139.5053 | 132.9707 | 0.0015011 | 0.00010019 | 0.95
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 13 | i13 | 139.5053 | 153.6403 | 146.5728 | 0.0015011 | 9.3497e-05 | 0.97
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 14 | i14 | 153.6403 | 169.3773 | 161.5088 | 0.0015011 | 8.9751e-05 | 0.99
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 15 | i15 | 169.3773 | 186.5858 | 177.9816 | 0.0015011 | 9.217e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 16 | i16 | 186.5858 | 205.4507 | 196.0183 | 0.0015011 | 8.6135e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 17 | i17 | 205.4507 | 227.054 | 216.2524 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 18 | i18 | 227.054 | 250.9369 | 238.9955 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 19 | i19 | 250.9369 | 275.8583 | 263.3976 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 20 | i20 | 275.8583 | 303.6288 | 289.7436 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 21 | i21 | 303.6288 | 334.1802 | 318.9045 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 22 | i22 | 334.1802 | 368.275 | 351.2276 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 23 | i23 | 368.275 | 405.8295 | 387.0522 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 24 | i24 | 405.8295 | 447.2136 | 426.5215 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 25 | i25 | 447.2136 | 492.8178 | 470.0157 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 26 | i26 | 492.8178 | 543.0725 | 517.9452 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 27 | i27 | 543.0725 | 598.4519 | 570.7622 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 28 | i28 | 598.4519 | 659.4785 | 628.9652 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 29 | i29 | 659.4785 | 726.7282 | 693.1033 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 30 | i30 | 726.7282 | 800.8357 | 763.782 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 31 | i31 | 800.8357 | 882.5003 | 841.668 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 32 | i32 | 882.5003 | 972.4925 | 927.4964 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 33 | i33 | 972.4925 | 1071.6616 | 1022.077 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 34 | i34 | 1071.6616 | 1180.9433 | 1126.3025 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 35 | i35 | 1180.9433 | 1301.369 | 1241.1562 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 36 | i36 | 1301.369 | 1444.1932 | 1372.7811 | 0.0015011 | 8.6202e-05 | 0.36667
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 37 | i37 | 1444.1932 | 1599.8498 | 1522.0215 | 0.0015011 | 8.6202e-05 | 0.8
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 38 | i38 | 1599.8498 | 1763.0124 | 1681.4311 | 0.0015011 | 8.6202e-05 | 0.75
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 39 | i39 | 1763.0124 | 1937.1886 | 1850.1005 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 40 | i40 | 1937.1886 | 2125.886 | 2031.5373 | 0.0015011 | 8.6202e-05 | 0.6
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 41 | i41 | 2125.886 | 2332.6122 | 2229.2491 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 42 | i42 | 2332.6122 | 2560.8748 | 2446.7435 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 43 | i43 | 2560.8748 | 2814.1814 | 2687.5281 | 0.0015011 | 8.6202e-05 | 0.952
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 44 | i44 | 2814.1814 | 3096.0396 | 2955.1105 | 0.0015011 | 8.6202e-05 | 0.87
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 45 | i45 | 3096.0396 | 3409.957 | 3252.9983 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 46 | i46 | 3409.957 | 3759.4412 | 3584.6991 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 47 | i47 | 3759.4412 | 4147.9998 | 3953.7205 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 48 | i48 | 4147.9998 | 4579.1404 | 4363.5701 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 49 | i49 | 4579.1404 | 5056.3706 | 4817.7555 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 50 | i50 | 5056.3706 | 5583.198 | 5319.7843 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 51 | i51 | 5583.198 | 6163.1302 | 5873.1641 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 52 | i52 | 6163.1302 | 6799.6748 | 6481.4025 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 53 | i53 | 6799.6748 | 7496.3394 | 7148.0071 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 54 | i54 | 7496.3394 | 8256.6316 | 7876.4855 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 55 | i55 | 8256.6316 | 9084.059 | 8670.3453 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 56 | i56 | 9084.059 | 9982.1292 | 9533.0941 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 57 | i57 | 9982.1292 | 10954.3498 | 10468.2395 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 58 | i58 | 10954.3498 | 12004.2284 | 11479.2891 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 59 | i59 | 12004.2284 | 13135.2726 | 12569.7505 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 60 | i60 | 13135.2726 | 14350.99 | 13743.1313 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 61 | i61 | 14350.99 | 15654.8882 | 15002.9391 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 62 | i62 | 15654.8882 | 17050.4748 | 16352.6815 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T0_Counts | 63 | i63 | 17050.4748 | 20000 | 18525.2374 | 0.0015011 | 8.6202e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 0 | i00 | 39.2625 | 43.9398 | 41.6011 | 0.0017478 | 0.00077038 | 0.00010417
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 1 | i01 | 43.9398 | 48.4219 | 46.1809 | 0.0017478 | 0.0028025 | 0.03
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 2 | i02 | 48.4219 | 52.7317 | 50.5768 | 0.0017478 | 0.00019859 | 0.3
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 3 | i03 | 52.7317 | 57.9075 | 55.3196 | 0.0017478 | 0.00013998 | 0.4
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 4 | i04 | 57.9075 | 64.8216 | 61.3645 | 0.0017478 | 0.00014511 | 0.43333
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 5 | i05 | 64.8216 | 70.4957 | 67.6587 | 0.0017478 | 0.00012479 | 0.7
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 6 | i06 | 70.4957 | 77.7766 | 74.1361 | 0.0017478 | 0.00012129 | 0.75
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 7 | i07 | 77.7766 | 85.6695 | 81.723 | 0.0017478 | 0.00012118 | 0.8
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 8 | i08 | 85.6695 | 94.1679 | 89.9187 | 0.0017478 | 0.00011901 | 0.82
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 9 | i09 | 94.1679 | 104.0937 | 99.1308 | 0.0017478 | 0.00012706 | 0.85
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 10 | i10 | 104.0937 | 114.5947 | 109.3442 | 0.0017478 | 0.00011648 | 0.9
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 11 | i11 | 114.5947 | 126.4362 | 120.5154 | 0.0017478 | 0.00011317 | 0.93
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 12 | i12 | 126.4362 | 139.5053 | 132.9707 | 0.0017478 | 0.00011195 | 0.95
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 13 | i13 | 139.5053 | 153.6403 | 146.5728 | 0.0017478 | 0.00010506 | 0.97
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 14 | i14 | 153.6403 | 169.3773 | 161.5088 | 0.0017478 | 0.00010968 | 0.99
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 15 | i15 | 169.3773 | 186.5858 | 177.9816 | 0.0017478 | 9.936e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 16 | i16 | 186.5858 | 205.4507 | 196.0183 | 0.0017478 | 0.00010153 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 17 | i17 | 205.4507 | 227.054 | 216.2524 | 0.0017478 | 0.00010002 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 18 | i18 | 227.054 | 250.9369 | 238.9955 | 0.0017478 | 0.00010257 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 19 | i19 | 250.9369 | 275.8583 | 263.3976 | 0.0017478 | 8.8299e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 20 | i20 | 275.8583 | 303.6288 | 289.7436 | 0.0017478 | 8.1931e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 21 | i21 | 303.6288 | 334.1802 | 318.9045 | 0.0017478 | 0.00010894 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 22 | i22 | 334.1802 | 368.275 | 351.2276 | 0.0017478 | 6.6837e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 23 | i23 | 368.275 | 405.8295 | 387.0522 | 0.0017478 | 9.0125e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 24 | i24 | 405.8295 | 447.2136 | 426.5215 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 25 | i25 | 447.2136 | 492.8178 | 470.0157 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 26 | i26 | 492.8178 | 543.0725 | 517.9452 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 27 | i27 | 543.0725 | 598.4519 | 570.7622 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 28 | i28 | 598.4519 | 659.4785 | 628.9652 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 29 | i29 | 659.4785 | 726.7282 | 693.1033 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 30 | i30 | 726.7282 | 800.8357 | 763.782 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 31 | i31 | 800.8357 | 882.5003 | 841.668 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 32 | i32 | 882.5003 | 972.4925 | 927.4964 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 33 | i33 | 972.4925 | 1071.6616 | 1022.077 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 34 | i34 | 1071.6616 | 1180.9433 | 1126.3025 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 35 | i35 | 1180.9433 | 1301.369 | 1241.1562 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 36 | i36 | 1301.369 | 1444.1932 | 1372.7811 | 0.0017478 | 9.3454e-05 | 0.36667
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 37 | i37 | 1444.1932 | 1599.8498 | 1522.0215 | 0.0017478 | 9.3454e-05 | 0.8
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 38 | i38 | 1599.8498 | 1763.0124 | 1681.4311 | 0.0017478 | 9.3454e-05 | 0.75
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 39 | i39 | 1763.0124 | 1937.1886 | 1850.1005 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 40 | i40 | 1937.1886 | 2125.886 | 2031.5373 | 0.0017478 | 9.3454e-05 | 0.6
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 41 | i41 | 2125.886 | 2332.6122 | 2229.2491 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 42 | i42 | 2332.6122 | 2560.8748 | 2446.7435 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 43 | i43 | 2560.8748 | 2814.1814 | 2687.5281 | 0.0017478 | 9.3454e-05 | 0.952
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 44 | i44 | 2814.1814 | 3096.0396 | 2955.1105 | 0.0017478 | 9.3454e-05 | 0.87
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 45 | i45 | 3096.0396 | 3409.957 | 3252.9983 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 46 | i46 | 3409.957 | 3759.4412 | 3584.6991 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 47 | i47 | 3759.4412 | 4147.9998 | 3953.7205 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 48 | i48 | 4147.9998 | 4579.1404 | 4363.5701 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 49 | i49 | 4579.1404 | 5056.3706 | 4817.7555 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 50 | i50 | 5056.3706 | 5583.198 | 5319.7843 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 51 | i51 | 5583.198 | 6163.1302 | 5873.1641 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 52 | i52 | 6163.1302 | 6799.6748 | 6481.4025 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 53 | i53 | 6799.6748 | 7496.3394 | 7148.0071 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 54 | i54 | 7496.3394 | 8256.6316 | 7876.4855 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 55 | i55 | 8256.6316 | 9084.059 | 8670.3453 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 56 | i56 | 9084.059 | 9982.1292 | 9533.0941 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 57 | i57 | 9982.1292 | 10954.3498 | 10468.2395 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 58 | i58 | 10954.3498 | 12004.2284 | 11479.2891 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 59 | i59 | 12004.2284 | 13135.2726 | 12569.7505 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 60 | i60 | 13135.2726 | 14350.99 | 13743.1313 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 61 | i61 | 14350.99 | 15654.8882 | 15002.9391 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 62 | i62 | 15654.8882 | 17050.4748 | 16352.6815 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T1_Counts | 63 | i63 | 17050.4748 | 20000 | 18525.2374 | 0.0017478 | 9.3454e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 0 | i00 | 39.2625 | 43.9398 | 41.6011 | 0.0016106 | 0.00047524 | 0.00010417
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 1 | i01 | 43.9398 | 48.4219 | 46.1809 | 0.0016106 | 0.0017867 | 0.03
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 2 | i02 | 48.4219 | 52.7317 | 50.5768 | 0.0016106 | 0.00027147 | 0.24
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 3 | i03 | 52.7317 | 57.9075 | 55.3196 | 0.0016106 | 0.00019014 | 0.4
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 4 | i04 | 57.9075 | 64.8216 | 61.3645 | 0.0016106 | 0.000119 | 0.43333
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 5 | i05 | 64.8216 | 70.4957 | 67.6587 | 0.0016106 | 0.00010037 | 0.7
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 6 | i06 | 70.4957 | 77.7766 | 74.1361 | 0.0016106 | 0.00010703 | 0.75
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 7 | i07 | 77.7766 | 85.6695 | 81.723 | 0.0016106 | 0.00010459 | 0.8
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 8 | i08 | 85.6695 | 94.1679 | 89.9187 | 0.0016106 | 0.00010537 | 0.82
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 9 | i09 | 94.1679 | 104.0937 | 99.1308 | 0.0016106 | 0.00010036 | 0.85
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 10 | i10 | 104.0937 | 114.5947 | 109.3442 | 0.0016106 | 9.5836e-05 | 0.9
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 11 | i11 | 114.5947 | 126.4362 | 120.5154 | 0.0016106 | 0.00010704 | 0.93
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 12 | i12 | 126.4362 | 139.5053 | 132.9707 | 0.0016106 | 0.00010344 | 0.95
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 13 | i13 | 139.5053 | 153.6403 | 146.5728 | 0.0016106 | 0.00010184 | 0.97
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 14 | i14 | 153.6403 | 169.3773 | 161.5088 | 0.0016106 | 9.9283e-05 | 0.99
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 15 | i15 | 169.3773 | 186.5858 | 177.9816 | 0.0016106 | 9.8448e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 16 | i16 | 186.5858 | 205.4507 | 196.0183 | 0.0016106 | 9.6827e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 17 | i17 | 205.4507 | 227.054 | 216.2524 | 0.0016106 | 9.241e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 18 | i18 | 227.054 | 250.9369 | 238.9955 | 0.0016106 | 8.7867e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 19 | i19 | 250.9369 | 275.8583 | 263.3976 | 0.0016106 | 9.3977e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 20 | i20 | 275.8583 | 303.6288 | 289.7436 | 0.0016106 | 7.906e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 21 | i21 | 303.6288 | 334.1802 | 318.9045 | 0.0016106 | 6.8117e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 22 | i22 | 334.1802 | 368.275 | 351.2276 | 0.0016106 | 7.549e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 23 | i23 | 368.275 | 405.8295 | 387.0522 | 0.0016106 | 9.355e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 24 | i24 | 405.8295 | 447.2136 | 426.5215 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 25 | i25 | 447.2136 | 492.8178 | 470.0157 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 26 | i26 | 492.8178 | 543.0725 | 517.9452 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 27 | i27 | 543.0725 | 598.4519 | 570.7622 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 28 | i28 | 598.4519 | 659.4785 | 628.9652 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 29 | i29 | 659.4785 | 726.7282 | 693.1033 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 30 | i30 | 726.7282 | 800.8357 | 763.782 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 31 | i31 | 800.8357 | 882.5003 | 841.668 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 32 | i32 | 882.5003 | 972.4925 | 927.4964 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 33 | i33 | 972.4925 | 1071.6616 | 1022.077 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 34 | i34 | 1071.6616 | 1180.9433 | 1126.3025 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 35 | i35 | 1180.9433 | 1301.369 | 1241.1562 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 36 | i36 | 1301.369 | 1444.1932 | 1372.7811 | 0.0016106 | 8.236e-05 | 0.36667
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 37 | i37 | 1444.1932 | 1599.8498 | 1522.0215 | 0.0016106 | 8.236e-05 | 0.8
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 38 | i38 | 1599.8498 | 1763.0124 | 1681.4311 | 0.0016106 | 8.236e-05 | 0.75
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 39 | i39 | 1763.0124 | 1937.1886 | 1850.1005 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 40 | i40 | 1937.1886 | 2125.886 | 2031.5373 | 0.0016106 | 8.236e-05 | 0.6
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 41 | i41 | 2125.886 | 2332.6122 | 2229.2491 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 42 | i42 | 2332.6122 | 2560.8748 | 2446.7435 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 43 | i43 | 2560.8748 | 2814.1814 | 2687.5281 | 0.0016106 | 8.236e-05 | 0.952
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 44 | i44 | 2814.1814 | 3096.0396 | 2955.1105 | 0.0016106 | 8.236e-05 | 0.87
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 45 | i45 | 3096.0396 | 3409.957 | 3252.9983 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 46 | i46 | 3409.957 | 3759.4412 | 3584.6991 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 47 | i47 | 3759.4412 | 4147.9998 | 3953.7205 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 48 | i48 | 4147.9998 | 4579.1404 | 4363.5701 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 49 | i49 | 4579.1404 | 5056.3706 | 4817.7555 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 50 | i50 | 5056.3706 | 5583.198 | 5319.7843 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 51 | i51 | 5583.198 | 6163.1302 | 5873.1641 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 52 | i52 | 6163.1302 | 6799.6748 | 6481.4025 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 53 | i53 | 6799.6748 | 7496.3394 | 7148.0071 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 54 | i54 | 7496.3394 | 8256.6316 | 7876.4855 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 55 | i55 | 8256.6316 | 9084.059 | 8670.3453 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 56 | i56 | 9084.059 | 9982.1292 | 9533.0941 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 57 | i57 | 9982.1292 | 10954.3498 | 10468.2395 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 58 | i58 | 10954.3498 | 12004.2284 | 11479.2891 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 59 | i59 | 12004.2284 | 13135.2726 | 12569.7505 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 60 | i60 | 13135.2726 | 14350.99 | 13743.1313 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 61 | i61 | 14350.99 | 15654.8882 | 15002.9391 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 62 | i62 | 15654.8882 | 17050.4748 | 16352.6815 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T2_Counts | 63 | i63 | 17050.4748 | 20000 | 18525.2374 | 0.0016106 | 8.236e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 0 | i00 | 39.2625 | 43.9398 | 41.6011 | 0.0011626 | 0.00072365 | 4.1667e-05
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 1 | i01 | 43.9398 | 48.4219 | 46.1809 | 0.0011626 | 0.0053377 | 0.006
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 2 | i02 | 48.4219 | 52.7317 | 50.5768 | 0.0011626 | 0.00017995 | 0.24
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 3 | i03 | 52.7317 | 57.9075 | 55.3196 | 0.0011626 | 0.00011247 | 0.4
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 4 | i04 | 57.9075 | 64.8216 | 61.3645 | 0.0011626 | 9.6491e-05 | 0.43333
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 5 | i05 | 64.8216 | 70.4957 | 67.6587 | 0.0011626 | 9.9356e-05 | 0.7
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 6 | i06 | 70.4957 | 77.7766 | 74.1361 | 0.0011626 | 7.8559e-05 | 0.75
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 7 | i07 | 77.7766 | 85.6695 | 81.723 | 0.0011626 | 7.6701e-05 | 0.8
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 8 | i08 | 85.6695 | 94.1679 | 89.9187 | 0.0011626 | 7.1454e-05 | 0.82
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 9 | i09 | 94.1679 | 104.0937 | 99.1308 | 0.0011626 | 7.6358e-05 | 0.85
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 10 | i10 | 104.0937 | 114.5947 | 109.3442 | 0.0011626 | 7.4047e-05 | 0.9
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 11 | i11 | 114.5947 | 126.4362 | 120.5154 | 0.0011626 | 7.4984e-05 | 0.93
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 12 | i12 | 126.4362 | 139.5053 | 132.9707 | 0.0011626 | 7.1779e-05 | 0.95
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 13 | i13 | 139.5053 | 153.6403 | 146.5728 | 0.0011626 | 7.7085e-05 | 0.97
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 14 | i14 | 153.6403 | 169.3773 | 161.5088 | 0.0011626 | 6.7316e-05 | 0.99
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 15 | i15 | 169.3773 | 186.5858 | 177.9816 | 0.0011626 | 8.5206e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 16 | i16 | 186.5858 | 205.4507 | 196.0183 | 0.0011626 | 7.2537e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 17 | i17 | 205.4507 | 227.054 | 216.2524 | 0.0011626 | 6.1671e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 18 | i18 | 227.054 | 250.9369 | 238.9955 | 0.0011626 | 7.7307e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 19 | i19 | 250.9369 | 275.8583 | 263.3976 | 0.0011626 | 7.7951e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 20 | i20 | 275.8583 | 303.6288 | 289.7436 | 0.0011626 | 7.1537e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 21 | i21 | 303.6288 | 334.1802 | 318.9045 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 22 | i22 | 334.1802 | 368.275 | 351.2276 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 23 | i23 | 368.275 | 405.8295 | 387.0522 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 24 | i24 | 405.8295 | 447.2136 | 426.5215 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 25 | i25 | 447.2136 | 492.8178 | 470.0157 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 26 | i26 | 492.8178 | 543.0725 | 517.9452 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 27 | i27 | 543.0725 | 598.4519 | 570.7622 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 28 | i28 | 598.4519 | 659.4785 | 628.9652 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 29 | i29 | 659.4785 | 726.7282 | 693.1033 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 30 | i30 | 726.7282 | 800.8357 | 763.782 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 31 | i31 | 800.8357 | 882.5003 | 841.668 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 32 | i32 | 882.5003 | 972.4925 | 927.4964 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 33 | i33 | 972.4925 | 1071.6616 | 1022.077 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 34 | i34 | 1071.6616 | 1180.9433 | 1126.3025 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 35 | i35 | 1180.9433 | 1301.369 | 1241.1562 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 36 | i36 | 1301.369 | 1444.1932 | 1372.7811 | 0.0011626 | 6.858e-05 | 0.36667
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 37 | i37 | 1444.1932 | 1599.8498 | 1522.0215 | 0.0011626 | 6.858e-05 | 0.8
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 38 | i38 | 1599.8498 | 1763.0124 | 1681.4311 | 0.0011626 | 6.858e-05 | 0.75
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 39 | i39 | 1763.0124 | 1937.1886 | 1850.1005 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 40 | i40 | 1937.1886 | 2125.886 | 2031.5373 | 0.0011626 | 6.858e-05 | 0.6
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 41 | i41 | 2125.886 | 2332.6122 | 2229.2491 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 42 | i42 | 2332.6122 | 2560.8748 | 2446.7435 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 43 | i43 | 2560.8748 | 2814.1814 | 2687.5281 | 0.0011626 | 6.858e-05 | 0.952
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 44 | i44 | 2814.1814 | 3096.0396 | 2955.1105 | 0.0011626 | 6.858e-05 | 0.87
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 45 | i45 | 3096.0396 | 3409.957 | 3252.9983 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 46 | i46 | 3409.957 | 3759.4412 | 3584.6991 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 47 | i47 | 3759.4412 | 4147.9998 | 3953.7205 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 48 | i48 | 4147.9998 | 4579.1404 | 4363.5701 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 49 | i49 | 4579.1404 | 5056.3706 | 4817.7555 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 50 | i50 | 5056.3706 | 5583.198 | 5319.7843 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 51 | i51 | 5583.198 | 6163.1302 | 5873.1641 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 52 | i52 | 6163.1302 | 6799.6748 | 6481.4025 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 53 | i53 | 6799.6748 | 7496.3394 | 7148.0071 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 54 | i54 | 7496.3394 | 8256.6316 | 7876.4855 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 55 | i55 | 8256.6316 | 9084.059 | 8670.3453 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 56 | i56 | 9084.059 | 9982.1292 | 9533.0941 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 57 | i57 | 9982.1292 | 10954.3498 | 10468.2395 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 58 | i58 | 10954.3498 | 12004.2284 | 11479.2891 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 59 | i59 | 12004.2284 | 13135.2726 | 12569.7505 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 60 | i60 | 13135.2726 | 14350.99 | 13743.1313 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 61 | i61 | 14350.99 | 15654.8882 | 15002.9391 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 62 | i62 | 15654.8882 | 17050.4748 | 16352.6815 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T3_Counts | 63 | i63 | 17050.4748 | 20000 | 18525.2374 | 0.0011626 | 6.858e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 0 | i00 | 39.2625 | 43.9398 | 41.6011 | 0.0015554 | 0.00093435 | 0.00083333
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 1 | i01 | 43.9398 | 48.4219 | 46.1809 | 0.0015554 | 0.00027632 | 0.18
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 2 | i02 | 48.4219 | 52.7317 | 50.5768 | 0.0015554 | 0.00014526 | 0.288
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 3 | i03 | 52.7317 | 57.9075 | 55.3196 | 0.0015554 | 0.00013662 | 0.4
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 4 | i04 | 57.9075 | 64.8216 | 61.3645 | 0.0015554 | 0.00011002 | 0.43333
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 5 | i05 | 64.8216 | 70.4957 | 67.6587 | 0.0015554 | 0.00010664 | 0.7
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 6 | i06 | 70.4957 | 77.7766 | 74.1361 | 0.0015554 | 0.0001169 | 0.75
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 7 | i07 | 77.7766 | 85.6695 | 81.723 | 0.0015554 | 0.00011174 | 0.8
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 8 | i08 | 85.6695 | 94.1679 | 89.9187 | 0.0015554 | 0.00010973 | 0.82
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 9 | i09 | 94.1679 | 104.0937 | 99.1308 | 0.0015554 | 0.00010727 | 0.85
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 10 | i10 | 104.0937 | 114.5947 | 109.3442 | 0.0015554 | 0.00010961 | 0.9
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 11 | i11 | 114.5947 | 126.4362 | 120.5154 | 0.0015554 | 0.00010572 | 0.93
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 12 | i12 | 126.4362 | 139.5053 | 132.9707 | 0.0015554 | 9.9862e-05 | 0.95
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 13 | i13 | 139.5053 | 153.6403 | 146.5728 | 0.0015554 | 0.00010273 | 0.97
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 14 | i14 | 153.6403 | 169.3773 | 161.5088 | 0.0015554 | 0.00010387 | 0.99
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 15 | i15 | 169.3773 | 186.5858 | 177.9816 | 0.0015554 | 0.00011857 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 16 | i16 | 186.5858 | 205.4507 | 196.0183 | 0.0015554 | 9.2993e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 17 | i17 | 205.4507 | 227.054 | 216.2524 | 0.0015554 | 7.481e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 18 | i18 | 227.054 | 250.9369 | 238.9955 | 0.0015554 | 0.00010242 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 19 | i19 | 250.9369 | 275.8583 | 263.3976 | 0.0015554 | 8.6003e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 20 | i20 | 275.8583 | 303.6288 | 289.7436 | 0.0015554 | 8.9996e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 21 | i21 | 303.6288 | 334.1802 | 318.9045 | 0.0015554 | 7.7917e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 22 | i22 | 334.1802 | 368.275 | 351.2276 | 0.0015554 | 7.9428e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 23 | i23 | 368.275 | 405.8295 | 387.0522 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 24 | i24 | 405.8295 | 447.2136 | 426.5215 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 25 | i25 | 447.2136 | 492.8178 | 470.0157 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 26 | i26 | 492.8178 | 543.0725 | 517.9452 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 27 | i27 | 543.0725 | 598.4519 | 570.7622 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 28 | i28 | 598.4519 | 659.4785 | 628.9652 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 29 | i29 | 659.4785 | 726.7282 | 693.1033 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 30 | i30 | 726.7282 | 800.8357 | 763.782 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 31 | i31 | 800.8357 | 882.5003 | 841.668 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 32 | i32 | 882.5003 | 972.4925 | 927.4964 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 33 | i33 | 972.4925 | 1071.6616 | 1022.077 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 34 | i34 | 1071.6616 | 1180.9433 | 1126.3025 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 35 | i35 | 1180.9433 | 1301.369 | 1241.1562 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 36 | i36 | 1301.369 | 1444.1932 | 1372.7811 | 0.0015554 | 8.4337e-05 | 0.36667
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 37 | i37 | 1444.1932 | 1599.8498 | 1522.0215 | 0.0015554 | 8.4337e-05 | 0.8
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 38 | i38 | 1599.8498 | 1763.0124 | 1681.4311 | 0.0015554 | 8.4337e-05 | 0.75
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 39 | i39 | 1763.0124 | 1937.1886 | 1850.1005 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 40 | i40 | 1937.1886 | 2125.886 | 2031.5373 | 0.0015554 | 8.4337e-05 | 0.6
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 41 | i41 | 2125.886 | 2332.6122 | 2229.2491 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 42 | i42 | 2332.6122 | 2560.8748 | 2446.7435 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 43 | i43 | 2560.8748 | 2814.1814 | 2687.5281 | 0.0015554 | 8.4337e-05 | 0.952
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 44 | i44 | 2814.1814 | 3096.0396 | 2955.1105 | 0.0015554 | 8.4337e-05 | 0.87
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 45 | i45 | 3096.0396 | 3409.957 | 3252.9983 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 46 | i46 | 3409.957 | 3759.4412 | 3584.6991 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 47 | i47 | 3759.4412 | 4147.9998 | 3953.7205 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 48 | i48 | 4147.9998 | 4579.1404 | 4363.5701 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 49 | i49 | 4579.1404 | 5056.3706 | 4817.7555 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 50 | i50 | 5056.3706 | 5583.198 | 5319.7843 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 51 | i51 | 5583.198 | 6163.1302 | 5873.1641 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 52 | i52 | 6163.1302 | 6799.6748 | 6481.4025 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 53 | i53 | 6799.6748 | 7496.3394 | 7148.0071 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 54 | i54 | 7496.3394 | 8256.6316 | 7876.4855 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 55 | i55 | 8256.6316 | 9084.059 | 8670.3453 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 56 | i56 | 9084.059 | 9982.1292 | 9533.0941 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 57 | i57 | 9982.1292 | 10954.3498 | 10468.2395 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 58 | i58 | 10954.3498 | 12004.2284 | 11479.2891 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 59 | i59 | 12004.2284 | 13135.2726 | 12569.7505 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 60 | i60 | 13135.2726 | 14350.99 | 13743.1313 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 61 | i61 | 14350.99 | 15654.8882 | 15002.9391 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 62 | i62 | 15654.8882 | 17050.4748 | 16352.6815 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T4_Counts | 63 | i63 | 17050.4748 | 20000 | 18525.2374 | 0.0015554 | 8.4337e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 0 | i00 | 39.2625 | 43.9398 | 41.6011 | 0.0012258 | 0.00069458 | 0.0014167
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 1 | i01 | 43.9398 | 48.4219 | 46.1809 | 0.0012258 | 0.00018711 | 0.18
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 2 | i02 | 48.4219 | 52.7317 | 50.5768 | 0.0012258 | 0.00011023 | 0.288
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 3 | i03 | 52.7317 | 57.9075 | 55.3196 | 0.0012258 | 9.9484e-05 | 0.4
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 4 | i04 | 57.9075 | 64.8216 | 61.3645 | 0.0012258 | 9.4095e-05 | 0.43333
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 5 | i05 | 64.8216 | 70.4957 | 67.6587 | 0.0012258 | 8.4944e-05 | 0.7
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 6 | i06 | 70.4957 | 77.7766 | 74.1361 | 0.0012258 | 8.2036e-05 | 0.75
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 7 | i07 | 77.7766 | 85.6695 | 81.723 | 0.0012258 | 8.7905e-05 | 0.8
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 8 | i08 | 85.6695 | 94.1679 | 89.9187 | 0.0012258 | 8.0312e-05 | 0.82
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 9 | i09 | 94.1679 | 104.0937 | 99.1308 | 0.0012258 | 7.2645e-05 | 0.85
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 10 | i10 | 104.0937 | 114.5947 | 109.3442 | 0.0012258 | 7.5654e-05 | 0.9
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 11 | i11 | 114.5947 | 126.4362 | 120.5154 | 0.0012258 | 7.9345e-05 | 0.93
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 12 | i12 | 126.4362 | 139.5053 | 132.9707 | 0.0012258 | 8.2781e-05 | 0.95
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 13 | i13 | 139.5053 | 153.6403 | 146.5728 | 0.0012258 | 7.9176e-05 | 0.97
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 14 | i14 | 153.6403 | 169.3773 | 161.5088 | 0.0012258 | 7.6424e-05 | 0.99
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 15 | i15 | 169.3773 | 186.5858 | 177.9816 | 0.0012258 | 7.2414e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 16 | i16 | 186.5858 | 205.4507 | 196.0183 | 0.0012258 | 7.3661e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 17 | i17 | 205.4507 | 227.054 | 216.2524 | 0.0012258 | 6.9973e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 18 | i18 | 227.054 | 250.9369 | 238.9955 | 0.0012258 | 6.6427e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 19 | i19 | 250.9369 | 275.8583 | 263.3976 | 0.0012258 | 5.587e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 20 | i20 | 275.8583 | 303.6288 | 289.7436 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 21 | i21 | 303.6288 | 334.1802 | 318.9045 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 22 | i22 | 334.1802 | 368.275 | 351.2276 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 23 | i23 | 368.275 | 405.8295 | 387.0522 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 24 | i24 | 405.8295 | 447.2136 | 426.5215 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 25 | i25 | 447.2136 | 492.8178 | 470.0157 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 26 | i26 | 492.8178 | 543.0725 | 517.9452 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 27 | i27 | 543.0725 | 598.4519 | 570.7622 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 28 | i28 | 598.4519 | 659.4785 | 628.9652 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 29 | i29 | 659.4785 | 726.7282 | 693.1033 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 30 | i30 | 726.7282 | 800.8357 | 763.782 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 31 | i31 | 800.8357 | 882.5003 | 841.668 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 32 | i32 | 882.5003 | 972.4925 | 927.4964 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 33 | i33 | 972.4925 | 1071.6616 | 1022.077 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 34 | i34 | 1071.6616 | 1180.9433 | 1126.3025 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 35 | i35 | 1180.9433 | 1301.369 | 1241.1562 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 36 | i36 | 1301.369 | 1444.1932 | 1372.7811 | 0.0012258 | 6.8182e-05 | 0.36667
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 37 | i37 | 1444.1932 | 1599.8498 | 1522.0215 | 0.0012258 | 6.8182e-05 | 0.8
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 38 | i38 | 1599.8498 | 1763.0124 | 1681.4311 | 0.0012258 | 6.8182e-05 | 0.75
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 39 | i39 | 1763.0124 | 1937.1886 | 1850.1005 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 40 | i40 | 1937.1886 | 2125.886 | 2031.5373 | 0.0012258 | 6.8182e-05 | 0.6
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 41 | i41 | 2125.886 | 2332.6122 | 2229.2491 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 42 | i42 | 2332.6122 | 2560.8748 | 2446.7435 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 43 | i43 | 2560.8748 | 2814.1814 | 2687.5281 | 0.0012258 | 6.8182e-05 | 0.952
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 44 | i44 | 2814.1814 | 3096.0396 | 2955.1105 | 0.0012258 | 6.8182e-05 | 0.87
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 45 | i45 | 3096.0396 | 3409.957 | 3252.9983 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 46 | i46 | 3409.957 | 3759.4412 | 3584.6991 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 47 | i47 | 3759.4412 | 4147.9998 | 3953.7205 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 48 | i48 | 4147.9998 | 4579.1404 | 4363.5701 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 49 | i49 | 4579.1404 | 5056.3706 | 4817.7555 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 50 | i50 | 5056.3706 | 5583.198 | 5319.7843 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 51 | i51 | 5583.198 | 6163.1302 | 5873.1641 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 52 | i52 | 6163.1302 | 6799.6748 | 6481.4025 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 53 | i53 | 6799.6748 | 7496.3394 | 7148.0071 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 54 | i54 | 7496.3394 | 8256.6316 | 7876.4855 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 55 | i55 | 8256.6316 | 9084.059 | 8670.3453 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 56 | i56 | 9084.059 | 9982.1292 | 9533.0941 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 57 | i57 | 9982.1292 | 10954.3498 | 10468.2395 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 58 | i58 | 10954.3498 | 12004.2284 | 11479.2891 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 59 | i59 | 12004.2284 | 13135.2726 | 12569.7505 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 60 | i60 | 13135.2726 | 14350.99 | 13743.1313 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 61 | i61 | 14350.99 | 15654.8882 | 15002.9391 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 62 | i62 | 15654.8882 | 17050.4748 | 16352.6815 | 0.0012258 | 6.8182e-05 | 1
Ion Species Rate - High Energy Res Low Time Res | ISR_HELT | 414734467.1853 | 662731267.1829 | Ion | T5_Counts | 63 | i63 | 17050.4748 | 20000 | 18525.2374 | 0.0012258 | 6.8182e-05 | 1

## TOFxE_H
Product Description | Product Name | StartET | StopET | Species | Field Name | Channel | ChName | E_Low (Kev) | E_Hi (Kev) | Mid_E(KeV) | G (cm^2 sr) | G-small | efficiency
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 0 | tep00 | 40.4514 | 48.9683 | 44.7099 | 0.0017196 | 0.0012142 | 0.0056696
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 1 | tep01 | 48.9683 | 60.4167 | 54.6925 | 0.0017196 | 0.00015785 | 0.11034
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 2 | tep02 | 60.4167 | 73.7162 | 67.0665 | 0.0017196 | 0.00012142 | 0.1701
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 3 | tep03 | 73.7162 | 89.5237 | 81.62 | 0.0017196 | 0.00012142 | 0.2082
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 4 | tep04 | 89.5237 | 109.3732 | 99.4485 | 0.0017196 | 0.00012142 | 0.24596
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 5 | tep05 | 109.3732 | 133.0186 | 121.1959 | 0.0017196 | 0.00012142 | 0.27327
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 6 | tep06 | 133.0186 | 162.6611 | 147.8398 | 0.0017196 | 0.00012142 | 0.26891
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 7 | tep07 | 162.6611 | 198.3912 | 180.5261 | 0.0017196 | 0.00012142 | 0.23941
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 8 | tep08 | 198.3912 | 242.5173 | 220.4543 | 0.0017196 | 0.00012142 | 0.21361
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 9 | tep09 | 242.5173 | 295.2375 | 268.8774 | 0.0017196 | 0.00012142 | 0.19387
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 10 | tep10 | 295.2375 | 359.8868 | 327.5622 | 0.0017196 | 0.00012142 | 0.17046
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 11 | tep11 | 359.8868 | 440.2896 | 400.0882 | 0.0017196 | 0.00012142 | 0.14663
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 12 | tep12 | 440.2896 | 537.5453 | 488.9174 | 0.0017196 | 0.00012142 | 0.12538
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T0_Counts | 13 | tep13 | 537.5453 | 657.6351 | 597.5902 | 0.0017196 | 0.00012142 | 0.1021
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 0 | tep00 | 40.4514 | 48.9683 | 44.7099 | 0.0017478 | 0.0017981 | 0.0043613
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 1 | tep01 | 48.9683 | 60.4167 | 54.6925 | 0.0017478 | 0.00017981 | 0.11034
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 2 | tep02 | 60.4167 | 73.7162 | 67.0665 | 0.0017478 | 0.00011988 | 0.1701
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 3 | tep03 | 73.7162 | 89.5237 | 81.62 | 0.0017478 | 0.00011988 | 0.2082
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 4 | tep04 | 89.5237 | 109.3732 | 99.4485 | 0.0017478 | 0.00011988 | 0.24596
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 5 | tep05 | 109.3732 | 133.0186 | 121.1959 | 0.0017478 | 0.00011988 | 0.27327
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 6 | tep06 | 133.0186 | 162.6611 | 147.8398 | 0.0017478 | 0.00011988 | 0.26891
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 7 | tep07 | 162.6611 | 198.3912 | 180.5261 | 0.0017478 | 0.00011988 | 0.23941
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 8 | tep08 | 198.3912 | 242.5173 | 220.4543 | 0.0017478 | 0.00011988 | 0.21361
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 9 | tep09 | 242.5173 | 295.2375 | 268.8774 | 0.0017478 | 0.00011988 | 0.19387
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 10 | tep10 | 295.2375 | 359.8868 | 327.5622 | 0.0017478 | 0.00011988 | 0.17046
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 11 | tep11 | 359.8868 | 440.2896 | 400.0882 | 0.0017478 | 0.00011988 | 0.14663
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 12 | tep12 | 440.2896 | 537.5453 | 488.9174 | 0.0017478 | 0.00011988 | 0.12538
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T1_Counts | 13 | tep13 | 537.5453 | 657.6351 | 597.5902 | 0.0017478 | 0.00011988 | 0.1021
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 0 | tep00 | 40.4514 | 48.9683 | 44.7099 | 0.0019402 | 0.0016955 | 0.0047974
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 1 | tep01 | 48.9683 | 60.4167 | 54.6925 | 0.0019402 | 0.00020868 | 0.088271
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 2 | tep02 | 60.4167 | 73.7162 | 67.0665 | 0.0019402 | 0.00013042 | 0.1701
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 3 | tep03 | 73.7162 | 89.5237 | 81.62 | 0.0019402 | 0.00013042 | 0.2082
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 4 | tep04 | 89.5237 | 109.3732 | 99.4485 | 0.0019402 | 0.00013042 | 0.24596
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 5 | tep05 | 109.3732 | 133.0186 | 121.1959 | 0.0019402 | 0.00013042 | 0.27327
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 6 | tep06 | 133.0186 | 162.6611 | 147.8398 | 0.0019402 | 0.00013042 | 0.26891
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 7 | tep07 | 162.6611 | 198.3912 | 180.5261 | 0.0019402 | 0.00013042 | 0.23941
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 8 | tep08 | 198.3912 | 242.5173 | 220.4543 | 0.0019402 | 0.00013042 | 0.21361
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 9 | tep09 | 242.5173 | 295.2375 | 268.8774 | 0.0019402 | 0.00013042 | 0.19387
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 10 | tep10 | 295.2375 | 359.8868 | 327.5622 | 0.0019402 | 0.00013042 | 0.17046
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 11 | tep11 | 359.8868 | 440.2896 | 400.0882 | 0.0019402 | 0.00013042 | 0.14663
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 12 | tep12 | 440.2896 | 537.5453 | 488.9174 | 0.0019402 | 0.00013042 | 0.12538
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T2_Counts | 13 | tep13 | 537.5453 | 657.6351 | 597.5902 | 0.0019402 | 0.00013042 | 0.1021
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 0 | tep00 | 40.4514 | 48.9683 | 44.7099 | 0.001252 | 0.0012201 | 0.004274
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 1 | tep01 | 48.9683 | 60.4167 | 54.6925 | 0.001252 | 0.00012816 | 0.088271
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 2 | tep02 | 60.4167 | 73.7162 | 67.0665 | 0.001252 | 8.5442e-05 | 0.1701
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 3 | tep03 | 73.7162 | 89.5237 | 81.62 | 0.001252 | 8.5442e-05 | 0.2082
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 4 | tep04 | 89.5237 | 109.3732 | 99.4485 | 0.001252 | 8.5442e-05 | 0.24596
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 5 | tep05 | 109.3732 | 133.0186 | 121.1959 | 0.001252 | 8.5442e-05 | 0.27327
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 6 | tep06 | 133.0186 | 162.6611 | 147.8398 | 0.001252 | 8.5442e-05 | 0.26891
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 7 | tep07 | 162.6611 | 198.3912 | 180.5261 | 0.001252 | 8.5442e-05 | 0.23941
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 8 | tep08 | 198.3912 | 242.5173 | 220.4543 | 0.001252 | 8.5442e-05 | 0.21361
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 9 | tep09 | 242.5173 | 295.2375 | 268.8774 | 0.001252 | 8.5442e-05 | 0.19387
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 10 | tep10 | 295.2375 | 359.8868 | 327.5622 | 0.001252 | 8.5442e-05 | 0.17046
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 11 | tep11 | 359.8868 | 440.2896 | 400.0882 | 0.001252 | 8.5442e-05 | 0.14663
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 12 | tep12 | 440.2896 | 537.5453 | 488.9174 | 0.001252 | 8.5442e-05 | 0.12538
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T3_Counts | 13 | tep13 | 537.5453 | 657.6351 | 597.5902 | 0.001252 | 8.5442e-05 | 0.1021
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 0 | tep00 | 40.4514 | 48.9683 | 44.7099 | 0.0017224 | 0.00080383 | 0.0063784
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 1 | tep01 | 48.9683 | 60.4167 | 54.6925 | 0.0017224 | 0.00014928 | 0.11034
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 2 | tep02 | 60.4167 | 73.7162 | 67.0665 | 0.0017224 | 0.00011483 | 0.1701
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 3 | tep03 | 73.7162 | 89.5237 | 81.62 | 0.0017224 | 0.00011483 | 0.2082
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 4 | tep04 | 89.5237 | 109.3732 | 99.4485 | 0.0017224 | 0.00011483 | 0.24596
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 5 | tep05 | 109.3732 | 133.0186 | 121.1959 | 0.0017224 | 0.00011483 | 0.27327
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 6 | tep06 | 133.0186 | 162.6611 | 147.8398 | 0.0017224 | 0.00011483 | 0.26891
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 7 | tep07 | 162.6611 | 198.3912 | 180.5261 | 0.0017224 | 0.00011483 | 0.23941
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 8 | tep08 | 198.3912 | 242.5173 | 220.4543 | 0.0017224 | 0.00011483 | 0.21361
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 9 | tep09 | 242.5173 | 295.2375 | 268.8774 | 0.0017224 | 0.00011483 | 0.19387
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 10 | tep10 | 295.2375 | 359.8868 | 327.5622 | 0.0017224 | 0.00011483 | 0.17046
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 11 | tep11 | 359.8868 | 440.2896 | 400.0882 | 0.0017224 | 0.00011483 | 0.14663
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 12 | tep12 | 440.2896 | 537.5453 | 488.9174 | 0.0017224 | 0.00011483 | 0.12538
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T4_Counts | 13 | tep13 | 537.5453 | 657.6351 | 597.5902 | 0.0017224 | 0.00011483 | 0.1021
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 0 | tep00 | 40.4514 | 48.9683 | 44.7099 | 0.0013412 | 0.00041277 | 0.0085045
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 1 | tep01 | 48.9683 | 60.4167 | 54.6925 | 0.0013412 | 0.00011007 | 0.11034
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 2 | tep02 | 60.4167 | 73.7162 | 67.0665 | 0.0013412 | 9.1728e-05 | 0.1701
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 3 | tep03 | 73.7162 | 89.5237 | 81.62 | 0.0013412 | 9.1728e-05 | 0.2082
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 4 | tep04 | 89.5237 | 109.3732 | 99.4485 | 0.0013412 | 9.1728e-05 | 0.24596
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 5 | tep05 | 109.3732 | 133.0186 | 121.1959 | 0.0013412 | 9.1728e-05 | 0.27327
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 6 | tep06 | 133.0186 | 162.6611 | 147.8398 | 0.0013412 | 9.1728e-05 | 0.26891
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 7 | tep07 | 162.6611 | 198.3912 | 180.5261 | 0.0013412 | 9.1728e-05 | 0.23941
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 8 | tep08 | 198.3912 | 242.5173 | 220.4543 | 0.0013412 | 9.1728e-05 | 0.21361
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 9 | tep09 | 242.5173 | 295.2375 | 268.8774 | 0.0013412 | 9.1728e-05 | 0.19387
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 10 | tep10 | 295.2375 | 359.8868 | 327.5622 | 0.0013412 | 9.1728e-05 | 0.17046
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 11 | tep11 | 359.8868 | 440.2896 | 400.0882 | 0.0013412 | 9.1728e-05 | 0.14663
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 12 | tep12 | 440.2896 | 537.5453 | 488.9174 | 0.0013412 | 9.1728e-05 | 0.12538
TOFxE Hydrogen Rates | TOFxE_H | 414734467.1853 | 662731267.1829 | H | T5_Counts | 13 | tep13 | 537.5453 | 657.6351 | 597.5902 | 0.0013412 | 9.1728e-05 | 0.1021

## TOF-Trigger_Ion E
Product Description | Product Name | StartET | StopET | Species | Field Name | Channel | ChName | E_Low (Kev) | E_Hi (Kev) | Mid_E(KeV) | G (cm^2 sr) | G-small | efficiency
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 0 | tei00 | 39.2625 | 43.9398 | 41.6011 | 0.0017196 | 0.00052974 | 0.0070074
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 1 | tei01 | 43.9398 | 48.4219 | 46.1809 | 0.0017196 | 0.00090446 | 0.01694
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 2 | tei02 | 48.4219 | 52.7317 | 50.5768 | 0.0017196 | 0.00016379 | 0.092309
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 3 | tei03 | 52.7317 | 57.9075 | 55.3196 | 0.0017196 | 0.00015069 | 0.13221
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 4 | tei04 | 57.9075 | 64.8216 | 61.3645 | 0.0017196 | 0.00012908 | 0.13125
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 5 | tei05 | 64.8216 | 70.4957 | 67.6587 | 0.0017196 | 0.00013097 | 0.20824
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 6 | tei06 | 70.4957 | 77.7766 | 74.1361 | 0.0017196 | 0.00011826 | 0.22947
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 7 | tei07 | 77.7766 | 85.6695 | 81.723 | 0.0017196 | 0.00011627 | 0.24791
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 8 | tei08 | 85.6695 | 94.1679 | 89.9187 | 0.0017196 | 0.00012617 | 0.26125
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 9 | tei09 | 94.1679 | 104.0937 | 99.1308 | 0.0017196 | 0.00011433 | 0.27542
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 10 | tei10 | 104.0937 | 114.5947 | 109.3442 | 0.0017196 | 0.00010833 | 0.28416
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 11 | tei11 | 114.5947 | 126.4362 | 120.5154 | 0.0017196 | 0.00010872 | 0.28932
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 12 | tei12 | 126.4362 | 139.5053 | 132.9707 | 0.0017196 | 0.00011477 | 0.27851
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 13 | tei13 | 139.5053 | 153.6403 | 146.5728 | 0.0017196 | 0.0001071 | 0.2691
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 14 | tei14 | 153.6403 | 169.3773 | 161.5088 | 0.0017196 | 0.00010281 | 0.25183
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 15 | tei15 | 169.3773 | 186.5858 | 177.9816 | 0.0017196 | 0.00010558 | 0.23974
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 16 | tei16 | 186.5858 | 205.4507 | 196.0183 | 0.0017196 | 9.8671e-05 | 0.22846
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 17 | tei17 | 205.4507 | 227.054 | 216.2524 | 0.0017196 | 9.8748e-05 | 0.21408
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 18 | tei18 | 227.054 | 250.9369 | 238.9955 | 0.0017196 | 9.8748e-05 | 0.20386
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 19 | tei19 | 250.9369 | 275.8583 | 263.3976 | 0.0017196 | 9.8748e-05 | 0.19442
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 20 | tei20 | 275.8583 | 303.6288 | 289.7436 | 0.0017196 | 9.8748e-05 | 0.1825
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 21 | tei21 | 303.6288 | 334.1802 | 318.9045 | 0.0017196 | 9.8748e-05 | 0.1712
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 22 | tei22 | 334.1802 | 368.275 | 351.2276 | 0.0017196 | 9.8748e-05 | 0.15765
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 23 | tei23 | 368.275 | 405.8295 | 387.0522 | 0.0017196 | 9.8748e-05 | 0.14756
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 24 | tei24 | 405.8295 | 447.2136 | 426.5215 | 0.0017196 | 9.8748e-05 | 0.13545
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 25 | tei25 | 447.2136 | 492.8178 | 470.0157 | 0.0017196 | 9.8748e-05 | 0.1265
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 26 | tei26 | 492.8178 | 543.0725 | 517.9452 | 0.0017196 | 9.8748e-05 | 0.11568
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 27 | tei27 | 543.0725 | 598.4519 | 570.7622 | 0.0017196 | 9.8748e-05 | 0.1059
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 28 | tei28 | 598.4519 | 659.4785 | 628.9652 | 0.0017196 | 9.8748e-05 | 0.09436
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 29 | tei29 | 659.4785 | 726.7282 | 693.1033 | 0.0017196 | 9.8748e-05 | 0.087687
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 30 | tei30 | 726.7282 | 800.8357 | 763.782 | 0.0017196 | 9.8748e-05 | 0.079259
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 31 | tei31 | 800.8357 | 882.5003 | 841.668 | 0.0017196 | 9.8748e-05 | 0.073274
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 32 | tei32 | 882.5003 | 972.4925 | 927.4964 | 0.0017196 | 9.8748e-05 | 0.065584
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 33 | tei33 | 972.4925 | 1071.6616 | 1022.077 | 0.0017196 | 9.8748e-05 | 0.058248
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 34 | tei34 | 1071.6616 | 1180.9433 | 1126.3025 | 0.0017196 | 9.8748e-05 | 0.051242
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 35 | tei35 | 1180.9433 | 1301.369 | 1241.1562 | 0.0017196 | 9.8748e-05 | 0.044545
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 36 | tei36 | 1301.369 | 1444.1932 | 1372.7811 | 0.0017196 | 9.8748e-05 | 0.03992
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 37 | tei37 | 1444.1932 | 1599.8498 | 1522.0215 | 0.0017196 | 9.8748e-05 | 0.037253
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 38 | tei38 | 1599.8498 | 1763.0124 | 1681.4311 | 0.0017196 | 9.8748e-05 | 0.032993
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 39 | tei39 | 1763.0124 | 1937.1886 | 1850.1005 | 0.0017196 | 9.8748e-05 | 0.028933
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 40 | tei40 | 1937.1886 | 2125.886 | 2031.5373 | 0.0017196 | 9.8748e-05 | 0.025042
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 41 | tei41 | 2125.886 | 2332.6122 | 2229.2491 | 0.0017196 | 9.8748e-05 | 0.021299
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 42 | tei42 | 2332.6122 | 2560.8748 | 2446.7435 | 0.0017196 | 9.8748e-05 | 0.019301
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 43 | tei43 | 2560.8748 | 2814.1814 | 2687.5281 | 0.0017196 | 9.8748e-05 | 0.017372
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 44 | tei44 | 2814.1814 | 3096.0396 | 2955.1105 | 0.0017196 | 9.8748e-05 | 0.015509
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 45 | tei45 | 3096.0396 | 3409.957 | 3252.9983 | 0.0017196 | 9.8748e-05 | 0.013708
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 46 | tei46 | 3409.957 | 3759.4412 | 3584.6991 | 0.0017196 | 9.8748e-05 | 0.011969
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 47 | tei47 | 3759.4412 | 4147.9998 | 3953.7205 | 0.0017196 | 9.8748e-05 | 0.010288
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 48 | tei48 | 4147.9998 | 4579.1404 | 4363.5701 | 0.0017196 | 9.8748e-05 | 0.0086651
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 49 | tei49 | 4579.1404 | 5056.3706 | 4817.7555 | 0.0017196 | 9.8748e-05 | 0.0073807
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 50 | tei50 | 5056.3706 | 5583.198 | 5319.7843 | 0.0017196 | 9.8748e-05 | 0.0062792
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 51 | tei51 | 5583.198 | 6163.1302 | 5873.1641 | 0.0017196 | 9.8748e-05 | 0.0057633
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 52 | tei52 | 6163.1302 | 6799.6748 | 6481.4025 | 0.0017196 | 9.8748e-05 | 0.0049947
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 53 | tei53 | 6799.6748 | 7496.3394 | 7148.0071 | 0.0017196 | 9.8748e-05 | 0.0042511
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 54 | tei54 | 7496.3394 | 8256.6316 | 7876.4855 | 0.0017196 | 9.8748e-05 | 0.0037928
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 55 | tei55 | 8256.6316 | 9084.059 | 8670.3453 | 0.0017196 | 9.8748e-05 | 0.0030914
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 56 | tei56 | 9084.059 | 9982.1292 | 9533.0941 | 0.0017196 | 9.8748e-05 | 0.0027921
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 57 | tei57 | 9982.1292 | 10954.3498 | 10468.2395 | 0.0017196 | 9.8748e-05 | 0.0023768
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 58 | tei58 | 10954.3498 | 12004.2284 | 11479.2891 | 0.0017196 | 9.8748e-05 | 0.0020971
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 59 | tei59 | 12004.2284 | 13135.2726 | 12569.7505 | 0.0017196 | 9.8748e-05 | 0.0017037
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 60 | tei60 | 13135.2726 | 14350.99 | 13743.1313 | 0.0017196 | 9.8748e-05 | 0.066054
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 61 | tei61 | 14350.99 | 15654.8882 | 15002.9391 | 0.0017196 | 9.8748e-05 | 0.060472
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 62 | tei62 | 15654.8882 | 17050.4748 | 16352.6815 | 0.0017196 | 9.8748e-05 | 0.053871
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T0_Counts | 63 | tei63 | 17050.4748 | 20000 | 18525.2374 | 0.0017196 | 9.8748e-05 | 0.047173
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 0 | tei00 | 39.2625 | 43.9398 | 41.6011 | 0.0017478 | 0.00077038 | 0.0087593
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 1 | tei01 | 43.9398 | 48.4219 | 46.1809 | 0.0017478 | 0.0028025 | 0.010164
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 2 | tei02 | 48.4219 | 52.7317 | 50.5768 | 0.0017478 | 0.00019859 | 0.092309
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 3 | tei03 | 52.7317 | 57.9075 | 55.3196 | 0.0017478 | 0.00013998 | 0.13221
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 4 | tei04 | 57.9075 | 64.8216 | 61.3645 | 0.0017478 | 0.00014511 | 0.13125
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 5 | tei05 | 64.8216 | 70.4957 | 67.6587 | 0.0017478 | 0.00012479 | 0.20824
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 6 | tei06 | 70.4957 | 77.7766 | 74.1361 | 0.0017478 | 0.00012129 | 0.22947
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 7 | tei07 | 77.7766 | 85.6695 | 81.723 | 0.0017478 | 0.00012118 | 0.24791
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 8 | tei08 | 85.6695 | 94.1679 | 89.9187 | 0.0017478 | 0.00011901 | 0.26125
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 9 | tei09 | 94.1679 | 104.0937 | 99.1308 | 0.0017478 | 0.00012706 | 0.27542
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 10 | tei10 | 104.0937 | 114.5947 | 109.3442 | 0.0017478 | 0.00011648 | 0.28416
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 11 | tei11 | 114.5947 | 126.4362 | 120.5154 | 0.0017478 | 0.00011317 | 0.28932
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 12 | tei12 | 126.4362 | 139.5053 | 132.9707 | 0.0017478 | 0.00011195 | 0.27851
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 13 | tei13 | 139.5053 | 153.6403 | 146.5728 | 0.0017478 | 0.00010506 | 0.2691
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 14 | tei14 | 153.6403 | 169.3773 | 161.5088 | 0.0017478 | 0.00010968 | 0.25183
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 15 | tei15 | 169.3773 | 186.5858 | 177.9816 | 0.0017478 | 9.936e-05 | 0.23974
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 16 | tei16 | 186.5858 | 205.4507 | 196.0183 | 0.0017478 | 0.00010153 | 0.22846
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 17 | tei17 | 205.4507 | 227.054 | 216.2524 | 0.0017478 | 0.00010002 | 0.21408
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 18 | tei18 | 227.054 | 250.9369 | 238.9955 | 0.0017478 | 0.00010257 | 0.20386
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 19 | tei19 | 250.9369 | 275.8583 | 263.3976 | 0.0017478 | 8.8299e-05 | 0.19442
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 20 | tei20 | 275.8583 | 303.6288 | 289.7436 | 0.0017478 | 8.1931e-05 | 0.1825
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 21 | tei21 | 303.6288 | 334.1802 | 318.9045 | 0.0017478 | 0.00010894 | 0.1712
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 22 | tei22 | 334.1802 | 368.275 | 351.2276 | 0.0017478 | 6.6837e-05 | 0.15765
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 23 | tei23 | 368.275 | 405.8295 | 387.0522 | 0.0017478 | 9.0125e-05 | 0.14756
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 24 | tei24 | 405.8295 | 447.2136 | 426.5215 | 0.0017478 | 9.3454e-05 | 0.13545
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 25 | tei25 | 447.2136 | 492.8178 | 470.0157 | 0.0017478 | 9.3454e-05 | 0.1265
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 26 | tei26 | 492.8178 | 543.0725 | 517.9452 | 0.0017478 | 9.3454e-05 | 0.11568
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 27 | tei27 | 543.0725 | 598.4519 | 570.7622 | 0.0017478 | 9.3454e-05 | 0.1059
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 28 | tei28 | 598.4519 | 659.4785 | 628.9652 | 0.0017478 | 9.3454e-05 | 0.09436
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 29 | tei29 | 659.4785 | 726.7282 | 693.1033 | 0.0017478 | 9.3454e-05 | 0.087687
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 30 | tei30 | 726.7282 | 800.8357 | 763.782 | 0.0017478 | 9.3454e-05 | 0.079259
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 31 | tei31 | 800.8357 | 882.5003 | 841.668 | 0.0017478 | 9.3454e-05 | 0.073274
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 32 | tei32 | 882.5003 | 972.4925 | 927.4964 | 0.0017478 | 9.3454e-05 | 0.065584
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 33 | tei33 | 972.4925 | 1071.6616 | 1022.077 | 0.0017478 | 9.3454e-05 | 0.058248
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 34 | tei34 | 1071.6616 | 1180.9433 | 1126.3025 | 0.0017478 | 9.3454e-05 | 0.051242
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 35 | tei35 | 1180.9433 | 1301.369 | 1241.1562 | 0.0017478 | 9.3454e-05 | 0.044545
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 36 | tei36 | 1301.369 | 1444.1932 | 1372.7811 | 0.0017478 | 9.3454e-05 | 0.03992
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 37 | tei37 | 1444.1932 | 1599.8498 | 1522.0215 | 0.0017478 | 9.3454e-05 | 0.037253
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 38 | tei38 | 1599.8498 | 1763.0124 | 1681.4311 | 0.0017478 | 9.3454e-05 | 0.032993
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 39 | tei39 | 1763.0124 | 1937.1886 | 1850.1005 | 0.0017478 | 9.3454e-05 | 0.028933
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 40 | tei40 | 1937.1886 | 2125.886 | 2031.5373 | 0.0017478 | 9.3454e-05 | 0.025042
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 41 | tei41 | 2125.886 | 2332.6122 | 2229.2491 | 0.0017478 | 9.3454e-05 | 0.021299
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 42 | tei42 | 2332.6122 | 2560.8748 | 2446.7435 | 0.0017478 | 9.3454e-05 | 0.019301
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 43 | tei43 | 2560.8748 | 2814.1814 | 2687.5281 | 0.0017478 | 9.3454e-05 | 0.017372
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 44 | tei44 | 2814.1814 | 3096.0396 | 2955.1105 | 0.0017478 | 9.3454e-05 | 0.015509
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 45 | tei45 | 3096.0396 | 3409.957 | 3252.9983 | 0.0017478 | 9.3454e-05 | 0.013708
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 46 | tei46 | 3409.957 | 3759.4412 | 3584.6991 | 0.0017478 | 9.3454e-05 | 0.011969
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 47 | tei47 | 3759.4412 | 4147.9998 | 3953.7205 | 0.0017478 | 9.3454e-05 | 0.010288
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 48 | tei48 | 4147.9998 | 4579.1404 | 4363.5701 | 0.0017478 | 9.3454e-05 | 0.0086651
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 49 | tei49 | 4579.1404 | 5056.3706 | 4817.7555 | 0.0017478 | 9.3454e-05 | 0.0073807
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 50 | tei50 | 5056.3706 | 5583.198 | 5319.7843 | 0.0017478 | 9.3454e-05 | 0.0062792
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 51 | tei51 | 5583.198 | 6163.1302 | 5873.1641 | 0.0017478 | 9.3454e-05 | 0.0057633
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 52 | tei52 | 6163.1302 | 6799.6748 | 6481.4025 | 0.0017478 | 9.3454e-05 | 0.0049947
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 53 | tei53 | 6799.6748 | 7496.3394 | 7148.0071 | 0.0017478 | 9.3454e-05 | 0.0042511
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 54 | tei54 | 7496.3394 | 8256.6316 | 7876.4855 | 0.0017478 | 9.3454e-05 | 0.0037928
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 55 | tei55 | 8256.6316 | 9084.059 | 8670.3453 | 0.0017478 | 9.3454e-05 | 0.0030914
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 56 | tei56 | 9084.059 | 9982.1292 | 9533.0941 | 0.0017478 | 9.3454e-05 | 0.0027921
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 57 | tei57 | 9982.1292 | 10954.3498 | 10468.2395 | 0.0017478 | 9.3454e-05 | 0.0023768
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 58 | tei58 | 10954.3498 | 12004.2284 | 11479.2891 | 0.0017478 | 9.3454e-05 | 0.0020971
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 59 | tei59 | 12004.2284 | 13135.2726 | 12569.7505 | 0.0017478 | 9.3454e-05 | 0.0017037
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 60 | tei60 | 13135.2726 | 14350.99 | 13743.1313 | 0.0017478 | 9.3454e-05 | 0.066054
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 61 | tei61 | 14350.99 | 15654.8882 | 15002.9391 | 0.0017478 | 9.3454e-05 | 0.060472
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 62 | tei62 | 15654.8882 | 17050.4748 | 16352.6815 | 0.0017478 | 9.3454e-05 | 0.053871
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T1_Counts | 63 | tei63 | 17050.4748 | 20000 | 18525.2374 | 0.0017478 | 9.3454e-05 | 0.047173
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 0 | tei00 | 39.2625 | 43.9398 | 41.6011 | 0.0019402 | 0.0005725 | 0.0087593
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 1 | tei01 | 43.9398 | 48.4219 | 46.1809 | 0.0019402 | 0.0021523 | 0.010164
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 2 | tei02 | 48.4219 | 52.7317 | 50.5768 | 0.0019402 | 0.00032702 | 0.066463
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 3 | tei03 | 52.7317 | 57.9075 | 55.3196 | 0.0019402 | 0.00022905 | 0.13221
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 4 | tei04 | 57.9075 | 64.8216 | 61.3645 | 0.0019402 | 0.00014335 | 0.13125
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 5 | tei05 | 64.8216 | 70.4957 | 67.6587 | 0.0019402 | 0.00012091 | 0.20824
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 6 | tei06 | 70.4957 | 77.7766 | 74.1361 | 0.0019402 | 0.00012893 | 0.22947
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 7 | tei07 | 77.7766 | 85.6695 | 81.723 | 0.0019402 | 0.00012599 | 0.24791
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 8 | tei08 | 85.6695 | 94.1679 | 89.9187 | 0.0019402 | 0.00012693 | 0.26125
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 9 | tei09 | 94.1679 | 104.0937 | 99.1308 | 0.0019402 | 0.0001209 | 0.27542
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 10 | tei10 | 104.0937 | 114.5947 | 109.3442 | 0.0019402 | 0.00011545 | 0.28416
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 11 | tei11 | 114.5947 | 126.4362 | 120.5154 | 0.0019402 | 0.00012894 | 0.28932
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 12 | tei12 | 126.4362 | 139.5053 | 132.9707 | 0.0019402 | 0.00012461 | 0.27851
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 13 | tei13 | 139.5053 | 153.6403 | 146.5728 | 0.0019402 | 0.00012269 | 0.2691
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 14 | tei14 | 153.6403 | 169.3773 | 161.5088 | 0.0019402 | 0.0001196 | 0.25183
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 15 | tei15 | 169.3773 | 186.5858 | 177.9816 | 0.0019402 | 0.0001186 | 0.23974
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 16 | tei16 | 186.5858 | 205.4507 | 196.0183 | 0.0019402 | 0.00011664 | 0.22846
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 17 | tei17 | 205.4507 | 227.054 | 216.2524 | 0.0019402 | 0.00011132 | 0.21408
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 18 | tei18 | 227.054 | 250.9369 | 238.9955 | 0.0019402 | 0.00010585 | 0.20386
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 19 | tei19 | 250.9369 | 275.8583 | 263.3976 | 0.0019402 | 0.00011321 | 0.19442
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 20 | tei20 | 275.8583 | 303.6288 | 289.7436 | 0.0019402 | 9.524e-05 | 0.1825
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 21 | tei21 | 303.6288 | 334.1802 | 318.9045 | 0.0019402 | 8.2057e-05 | 0.1712
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 22 | tei22 | 334.1802 | 368.275 | 351.2276 | 0.0019402 | 9.0939e-05 | 0.15765
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 23 | tei23 | 368.275 | 405.8295 | 387.0522 | 0.0019402 | 0.00011269 | 0.14756
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 24 | tei24 | 405.8295 | 447.2136 | 426.5215 | 0.0019402 | 9.9215e-05 | 0.13545
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 25 | tei25 | 447.2136 | 492.8178 | 470.0157 | 0.0019402 | 9.9215e-05 | 0.1265
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 26 | tei26 | 492.8178 | 543.0725 | 517.9452 | 0.0019402 | 9.9215e-05 | 0.11568
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 27 | tei27 | 543.0725 | 598.4519 | 570.7622 | 0.0019402 | 9.9215e-05 | 0.1059
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 28 | tei28 | 598.4519 | 659.4785 | 628.9652 | 0.0019402 | 9.9215e-05 | 0.09436
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 29 | tei29 | 659.4785 | 726.7282 | 693.1033 | 0.0019402 | 9.9215e-05 | 0.087687
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 30 | tei30 | 726.7282 | 800.8357 | 763.782 | 0.0019402 | 9.9215e-05 | 0.079259
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 31 | tei31 | 800.8357 | 882.5003 | 841.668 | 0.0019402 | 9.9215e-05 | 0.073274
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 32 | tei32 | 882.5003 | 972.4925 | 927.4964 | 0.0019402 | 9.9215e-05 | 0.065584
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 33 | tei33 | 972.4925 | 1071.6616 | 1022.077 | 0.0019402 | 9.9215e-05 | 0.058248
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 34 | tei34 | 1071.6616 | 1180.9433 | 1126.3025 | 0.0019402 | 9.9215e-05 | 0.051242
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 35 | tei35 | 1180.9433 | 1301.369 | 1241.1562 | 0.0019402 | 9.9215e-05 | 0.044545
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 36 | tei36 | 1301.369 | 1444.1932 | 1372.7811 | 0.0019402 | 9.9215e-05 | 0.03992
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 37 | tei37 | 1444.1932 | 1599.8498 | 1522.0215 | 0.0019402 | 9.9215e-05 | 0.037253
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 38 | tei38 | 1599.8498 | 1763.0124 | 1681.4311 | 0.0019402 | 9.9215e-05 | 0.032993
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 39 | tei39 | 1763.0124 | 1937.1886 | 1850.1005 | 0.0019402 | 9.9215e-05 | 0.028933
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 40 | tei40 | 1937.1886 | 2125.886 | 2031.5373 | 0.0019402 | 9.9215e-05 | 0.025042
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 41 | tei41 | 2125.886 | 2332.6122 | 2229.2491 | 0.0019402 | 9.9215e-05 | 0.021299
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 42 | tei42 | 2332.6122 | 2560.8748 | 2446.7435 | 0.0019402 | 9.9215e-05 | 0.019301
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 43 | tei43 | 2560.8748 | 2814.1814 | 2687.5281 | 0.0019402 | 9.9215e-05 | 0.017372
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 44 | tei44 | 2814.1814 | 3096.0396 | 2955.1105 | 0.0019402 | 9.9215e-05 | 0.015509
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 45 | tei45 | 3096.0396 | 3409.957 | 3252.9983 | 0.0019402 | 9.9215e-05 | 0.013708
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 46 | tei46 | 3409.957 | 3759.4412 | 3584.6991 | 0.0019402 | 9.9215e-05 | 0.011969
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 47 | tei47 | 3759.4412 | 4147.9998 | 3953.7205 | 0.0019402 | 9.9215e-05 | 0.010288
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 48 | tei48 | 4147.9998 | 4579.1404 | 4363.5701 | 0.0019402 | 9.9215e-05 | 0.0086651
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 49 | tei49 | 4579.1404 | 5056.3706 | 4817.7555 | 0.0019402 | 9.9215e-05 | 0.0073807
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 50 | tei50 | 5056.3706 | 5583.198 | 5319.7843 | 0.0019402 | 9.9215e-05 | 0.0062792
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 51 | tei51 | 5583.198 | 6163.1302 | 5873.1641 | 0.0019402 | 9.9215e-05 | 0.0057633
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 52 | tei52 | 6163.1302 | 6799.6748 | 6481.4025 | 0.0019402 | 9.9215e-05 | 0.0049947
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 53 | tei53 | 6799.6748 | 7496.3394 | 7148.0071 | 0.0019402 | 9.9215e-05 | 0.0042511
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 54 | tei54 | 7496.3394 | 8256.6316 | 7876.4855 | 0.0019402 | 9.9215e-05 | 0.0037928
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 55 | tei55 | 8256.6316 | 9084.059 | 8670.3453 | 0.0019402 | 9.9215e-05 | 0.0030914
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 56 | tei56 | 9084.059 | 9982.1292 | 9533.0941 | 0.0019402 | 9.9215e-05 | 0.0027921
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 57 | tei57 | 9982.1292 | 10954.3498 | 10468.2395 | 0.0019402 | 9.9215e-05 | 0.0023768
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 58 | tei58 | 10954.3498 | 12004.2284 | 11479.2891 | 0.0019402 | 9.9215e-05 | 0.0020971
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 59 | tei59 | 12004.2284 | 13135.2726 | 12569.7505 | 0.0019402 | 9.9215e-05 | 0.0017037
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 60 | tei60 | 13135.2726 | 14350.99 | 13743.1313 | 0.0019402 | 9.9215e-05 | 0.066054
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 61 | tei61 | 14350.99 | 15654.8882 | 15002.9391 | 0.0019402 | 9.9215e-05 | 0.060472
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 62 | tei62 | 15654.8882 | 17050.4748 | 16352.6815 | 0.0019402 | 9.9215e-05 | 0.053871
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T2_Counts | 63 | tei63 | 17050.4748 | 20000 | 18525.2374 | 0.0019402 | 9.9215e-05 | 0.047173
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 0 | tei00 | 39.2625 | 43.9398 | 41.6011 | 0.001252 | 0.00077929 | 0.0046716
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 1 | tei01 | 43.9398 | 48.4219 | 46.1809 | 0.001252 | 0.0057481 | 0.0021174
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 2 | tei02 | 48.4219 | 52.7317 | 50.5768 | 0.001252 | 0.00019378 | 0.066463
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 3 | tei03 | 52.7317 | 57.9075 | 55.3196 | 0.001252 | 0.00012112 | 0.13221
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 4 | tei04 | 57.9075 | 64.8216 | 61.3645 | 0.001252 | 0.00010391 | 0.13125
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 5 | tei05 | 64.8216 | 70.4957 | 67.6587 | 0.001252 | 0.000107 | 0.20824
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 6 | tei06 | 70.4957 | 77.7766 | 74.1361 | 0.001252 | 8.4599e-05 | 0.22947
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 7 | tei07 | 77.7766 | 85.6695 | 81.723 | 0.001252 | 8.2599e-05 | 0.24791
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 8 | tei08 | 85.6695 | 94.1679 | 89.9187 | 0.001252 | 7.6948e-05 | 0.26125
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 9 | tei09 | 94.1679 | 104.0937 | 99.1308 | 0.001252 | 8.223e-05 | 0.27542
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 10 | tei10 | 104.0937 | 114.5947 | 109.3442 | 0.001252 | 7.9741e-05 | 0.28416
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 11 | tei11 | 114.5947 | 126.4362 | 120.5154 | 0.001252 | 8.0749e-05 | 0.28932
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 12 | tei12 | 126.4362 | 139.5053 | 132.9707 | 0.001252 | 7.7299e-05 | 0.27851
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 13 | tei13 | 139.5053 | 153.6403 | 146.5728 | 0.001252 | 8.3012e-05 | 0.2691
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 14 | tei14 | 153.6403 | 169.3773 | 161.5088 | 0.001252 | 7.2492e-05 | 0.25183
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 15 | tei15 | 169.3773 | 186.5858 | 177.9816 | 0.001252 | 9.1758e-05 | 0.23974
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 16 | tei16 | 186.5858 | 205.4507 | 196.0183 | 0.001252 | 7.8114e-05 | 0.22846
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 17 | tei17 | 205.4507 | 227.054 | 216.2524 | 0.001252 | 6.6413e-05 | 0.21408
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 18 | tei18 | 227.054 | 250.9369 | 238.9955 | 0.001252 | 8.3251e-05 | 0.20386
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 19 | tei19 | 250.9369 | 275.8583 | 263.3976 | 0.001252 | 8.3945e-05 | 0.19442
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 20 | tei20 | 275.8583 | 303.6288 | 289.7436 | 0.001252 | 7.7038e-05 | 0.1825
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 21 | tei21 | 303.6288 | 334.1802 | 318.9045 | 0.001252 | 7.3854e-05 | 0.1712
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 22 | tei22 | 334.1802 | 368.275 | 351.2276 | 0.001252 | 7.3854e-05 | 0.15765
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 23 | tei23 | 368.275 | 405.8295 | 387.0522 | 0.001252 | 7.3854e-05 | 0.14756
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 24 | tei24 | 405.8295 | 447.2136 | 426.5215 | 0.001252 | 7.3854e-05 | 0.13545
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 25 | tei25 | 447.2136 | 492.8178 | 470.0157 | 0.001252 | 7.3854e-05 | 0.1265
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 26 | tei26 | 492.8178 | 543.0725 | 517.9452 | 0.001252 | 7.3854e-05 | 0.11568
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 27 | tei27 | 543.0725 | 598.4519 | 570.7622 | 0.001252 | 7.3854e-05 | 0.1059
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 28 | tei28 | 598.4519 | 659.4785 | 628.9652 | 0.001252 | 7.3854e-05 | 0.09436
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 29 | tei29 | 659.4785 | 726.7282 | 693.1033 | 0.001252 | 7.3854e-05 | 0.087687
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 30 | tei30 | 726.7282 | 800.8357 | 763.782 | 0.001252 | 7.3854e-05 | 0.079259
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 31 | tei31 | 800.8357 | 882.5003 | 841.668 | 0.001252 | 7.3854e-05 | 0.073274
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 32 | tei32 | 882.5003 | 972.4925 | 927.4964 | 0.001252 | 7.3854e-05 | 0.065584
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 33 | tei33 | 972.4925 | 1071.6616 | 1022.077 | 0.001252 | 7.3854e-05 | 0.058248
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 34 | tei34 | 1071.6616 | 1180.9433 | 1126.3025 | 0.001252 | 7.3854e-05 | 0.051242
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 35 | tei35 | 1180.9433 | 1301.369 | 1241.1562 | 0.001252 | 7.3854e-05 | 0.044545
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 36 | tei36 | 1301.369 | 1444.1932 | 1372.7811 | 0.001252 | 7.3854e-05 | 0.03992
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 37 | tei37 | 1444.1932 | 1599.8498 | 1522.0215 | 0.001252 | 7.3854e-05 | 0.037253
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 38 | tei38 | 1599.8498 | 1763.0124 | 1681.4311 | 0.001252 | 7.3854e-05 | 0.032993
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 39 | tei39 | 1763.0124 | 1937.1886 | 1850.1005 | 0.001252 | 7.3854e-05 | 0.028933
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 40 | tei40 | 1937.1886 | 2125.886 | 2031.5373 | 0.001252 | 7.3854e-05 | 0.025042
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 41 | tei41 | 2125.886 | 2332.6122 | 2229.2491 | 0.001252 | 7.3854e-05 | 0.021299
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 42 | tei42 | 2332.6122 | 2560.8748 | 2446.7435 | 0.001252 | 7.3854e-05 | 0.019301
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 43 | tei43 | 2560.8748 | 2814.1814 | 2687.5281 | 0.001252 | 7.3854e-05 | 0.017372
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 44 | tei44 | 2814.1814 | 3096.0396 | 2955.1105 | 0.001252 | 7.3854e-05 | 0.015509
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 45 | tei45 | 3096.0396 | 3409.957 | 3252.9983 | 0.001252 | 7.3854e-05 | 0.013708
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 46 | tei46 | 3409.957 | 3759.4412 | 3584.6991 | 0.001252 | 7.3854e-05 | 0.011969
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 47 | tei47 | 3759.4412 | 4147.9998 | 3953.7205 | 0.001252 | 7.3854e-05 | 0.010288
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 48 | tei48 | 4147.9998 | 4579.1404 | 4363.5701 | 0.001252 | 7.3854e-05 | 0.0086651
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 49 | tei49 | 4579.1404 | 5056.3706 | 4817.7555 | 0.001252 | 7.3854e-05 | 0.0073807
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 50 | tei50 | 5056.3706 | 5583.198 | 5319.7843 | 0.001252 | 7.3854e-05 | 0.0062792
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 51 | tei51 | 5583.198 | 6163.1302 | 5873.1641 | 0.001252 | 7.3854e-05 | 0.0057633
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 52 | tei52 | 6163.1302 | 6799.6748 | 6481.4025 | 0.001252 | 7.3854e-05 | 0.0049947
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 53 | tei53 | 6799.6748 | 7496.3394 | 7148.0071 | 0.001252 | 7.3854e-05 | 0.0042511
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 54 | tei54 | 7496.3394 | 8256.6316 | 7876.4855 | 0.001252 | 7.3854e-05 | 0.0037928
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 55 | tei55 | 8256.6316 | 9084.059 | 8670.3453 | 0.001252 | 7.3854e-05 | 0.0030914
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 56 | tei56 | 9084.059 | 9982.1292 | 9533.0941 | 0.001252 | 7.3854e-05 | 0.0027921
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 57 | tei57 | 9982.1292 | 10954.3498 | 10468.2395 | 0.001252 | 7.3854e-05 | 0.0023768
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 58 | tei58 | 10954.3498 | 12004.2284 | 11479.2891 | 0.001252 | 7.3854e-05 | 0.0020971
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 59 | tei59 | 12004.2284 | 13135.2726 | 12569.7505 | 0.001252 | 7.3854e-05 | 0.0017037
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 60 | tei60 | 13135.2726 | 14350.99 | 13743.1313 | 0.001252 | 7.3854e-05 | 0.066054
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 61 | tei61 | 14350.99 | 15654.8882 | 15002.9391 | 0.001252 | 7.3854e-05 | 0.060472
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 62 | tei62 | 15654.8882 | 17050.4748 | 16352.6815 | 0.001252 | 7.3854e-05 | 0.053871
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T3_Counts | 63 | tei63 | 17050.4748 | 20000 | 18525.2374 | 0.001252 | 7.3854e-05 | 0.047173
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 0 | tei00 | 39.2625 | 43.9398 | 41.6011 | 0.0017224 | 0.0010346 | 0.0070074
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 1 | tei01 | 43.9398 | 48.4219 | 46.1809 | 0.0017224 | 0.00030597 | 0.056465
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 2 | tei02 | 48.4219 | 52.7317 | 50.5768 | 0.0017224 | 0.00016084 | 0.092309
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 3 | tei03 | 52.7317 | 57.9075 | 55.3196 | 0.0017224 | 0.00015129 | 0.13221
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 4 | tei04 | 57.9075 | 64.8216 | 61.3645 | 0.0017224 | 0.00012182 | 0.13125
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 5 | tei05 | 64.8216 | 70.4957 | 67.6587 | 0.0017224 | 0.00011808 | 0.20824
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 6 | tei06 | 70.4957 | 77.7766 | 74.1361 | 0.0017224 | 0.00012945 | 0.22947
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 7 | tei07 | 77.7766 | 85.6695 | 81.723 | 0.0017224 | 0.00012374 | 0.24791
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 8 | tei08 | 85.6695 | 94.1679 | 89.9187 | 0.0017224 | 0.00012151 | 0.26125
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 9 | tei09 | 94.1679 | 104.0937 | 99.1308 | 0.0017224 | 0.00011878 | 0.27542
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 10 | tei10 | 104.0937 | 114.5947 | 109.3442 | 0.0017224 | 0.00012138 | 0.28416
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 11 | tei11 | 114.5947 | 126.4362 | 120.5154 | 0.0017224 | 0.00011707 | 0.28932
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 12 | tei12 | 126.4362 | 139.5053 | 132.9707 | 0.0017224 | 0.00011058 | 0.27851
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 13 | tei13 | 139.5053 | 153.6403 | 146.5728 | 0.0017224 | 0.00011376 | 0.2691
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 14 | tei14 | 153.6403 | 169.3773 | 161.5088 | 0.0017224 | 0.00011502 | 0.25183
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 15 | tei15 | 169.3773 | 186.5858 | 177.9816 | 0.0017224 | 0.00013129 | 0.23974
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 16 | tei16 | 186.5858 | 205.4507 | 196.0183 | 0.0017224 | 0.00010297 | 0.22846
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 17 | tei17 | 205.4507 | 227.054 | 216.2524 | 0.0017224 | 8.2839e-05 | 0.21408
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 18 | tei18 | 227.054 | 250.9369 | 238.9955 | 0.0017224 | 0.00011341 | 0.20386
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 19 | tei19 | 250.9369 | 275.8583 | 263.3976 | 0.0017224 | 9.5233e-05 | 0.19442
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 20 | tei20 | 275.8583 | 303.6288 | 289.7436 | 0.0017224 | 9.9655e-05 | 0.1825
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 21 | tei21 | 303.6288 | 334.1802 | 318.9045 | 0.0017224 | 8.6279e-05 | 0.1712
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 22 | tei22 | 334.1802 | 368.275 | 351.2276 | 0.0017224 | 8.7953e-05 | 0.15765
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 23 | tei23 | 368.275 | 405.8295 | 387.0522 | 0.0017224 | 9.3388e-05 | 0.14756
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 24 | tei24 | 405.8295 | 447.2136 | 426.5215 | 0.0017224 | 9.3388e-05 | 0.13545
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 25 | tei25 | 447.2136 | 492.8178 | 470.0157 | 0.0017224 | 9.3388e-05 | 0.1265
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 26 | tei26 | 492.8178 | 543.0725 | 517.9452 | 0.0017224 | 9.3388e-05 | 0.11568
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 27 | tei27 | 543.0725 | 598.4519 | 570.7622 | 0.0017224 | 9.3388e-05 | 0.1059
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 28 | tei28 | 598.4519 | 659.4785 | 628.9652 | 0.0017224 | 9.3388e-05 | 0.09436
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 29 | tei29 | 659.4785 | 726.7282 | 693.1033 | 0.0017224 | 9.3388e-05 | 0.087687
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 30 | tei30 | 726.7282 | 800.8357 | 763.782 | 0.0017224 | 9.3388e-05 | 0.079259
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 31 | tei31 | 800.8357 | 882.5003 | 841.668 | 0.0017224 | 9.3388e-05 | 0.073274
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 32 | tei32 | 882.5003 | 972.4925 | 927.4964 | 0.0017224 | 9.3388e-05 | 0.065584
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 33 | tei33 | 972.4925 | 1071.6616 | 1022.077 | 0.0017224 | 9.3388e-05 | 0.058248
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 34 | tei34 | 1071.6616 | 1180.9433 | 1126.3025 | 0.0017224 | 9.3388e-05 | 0.051242
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 35 | tei35 | 1180.9433 | 1301.369 | 1241.1562 | 0.0017224 | 9.3388e-05 | 0.044545
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 36 | tei36 | 1301.369 | 1444.1932 | 1372.7811 | 0.0017224 | 9.3388e-05 | 0.03992
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 37 | tei37 | 1444.1932 | 1599.8498 | 1522.0215 | 0.0017224 | 9.3388e-05 | 0.037253
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 38 | tei38 | 1599.8498 | 1763.0124 | 1681.4311 | 0.0017224 | 9.3388e-05 | 0.032993
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 39 | tei39 | 1763.0124 | 1937.1886 | 1850.1005 | 0.0017224 | 9.3388e-05 | 0.028933
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 40 | tei40 | 1937.1886 | 2125.886 | 2031.5373 | 0.0017224 | 9.3388e-05 | 0.025042
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 41 | tei41 | 2125.886 | 2332.6122 | 2229.2491 | 0.0017224 | 9.3388e-05 | 0.021299
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 42 | tei42 | 2332.6122 | 2560.8748 | 2446.7435 | 0.0017224 | 9.3388e-05 | 0.019301
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 43 | tei43 | 2560.8748 | 2814.1814 | 2687.5281 | 0.0017224 | 9.3388e-05 | 0.017372
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 44 | tei44 | 2814.1814 | 3096.0396 | 2955.1105 | 0.0017224 | 9.3388e-05 | 0.015509
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 45 | tei45 | 3096.0396 | 3409.957 | 3252.9983 | 0.0017224 | 9.3388e-05 | 0.013708
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 46 | tei46 | 3409.957 | 3759.4412 | 3584.6991 | 0.0017224 | 9.3388e-05 | 0.011969
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 47 | tei47 | 3759.4412 | 4147.9998 | 3953.7205 | 0.0017224 | 9.3388e-05 | 0.010288
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 48 | tei48 | 4147.9998 | 4579.1404 | 4363.5701 | 0.0017224 | 9.3388e-05 | 0.0086651
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 49 | tei49 | 4579.1404 | 5056.3706 | 4817.7555 | 0.0017224 | 9.3388e-05 | 0.0073807
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 50 | tei50 | 5056.3706 | 5583.198 | 5319.7843 | 0.0017224 | 9.3388e-05 | 0.0062792
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 51 | tei51 | 5583.198 | 6163.1302 | 5873.1641 | 0.0017224 | 9.3388e-05 | 0.0057633
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 52 | tei52 | 6163.1302 | 6799.6748 | 6481.4025 | 0.0017224 | 9.3388e-05 | 0.0049947
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 53 | tei53 | 6799.6748 | 7496.3394 | 7148.0071 | 0.0017224 | 9.3388e-05 | 0.0042511
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 54 | tei54 | 7496.3394 | 8256.6316 | 7876.4855 | 0.0017224 | 9.3388e-05 | 0.0037928
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 55 | tei55 | 8256.6316 | 9084.059 | 8670.3453 | 0.0017224 | 9.3388e-05 | 0.0030914
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 56 | tei56 | 9084.059 | 9982.1292 | 9533.0941 | 0.0017224 | 9.3388e-05 | 0.0027921
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 57 | tei57 | 9982.1292 | 10954.3498 | 10468.2395 | 0.0017224 | 9.3388e-05 | 0.0023768
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 58 | tei58 | 10954.3498 | 12004.2284 | 11479.2891 | 0.0017224 | 9.3388e-05 | 0.0020971
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 59 | tei59 | 12004.2284 | 13135.2726 | 12569.7505 | 0.0017224 | 9.3388e-05 | 0.0017037
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 60 | tei60 | 13135.2726 | 14350.99 | 13743.1313 | 0.0017224 | 9.3388e-05 | 0.066054
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 61 | tei61 | 14350.99 | 15654.8882 | 15002.9391 | 0.0017224 | 9.3388e-05 | 0.060472
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 62 | tei62 | 15654.8882 | 17050.4748 | 16352.6815 | 0.0017224 | 9.3388e-05 | 0.053871
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T4_Counts | 63 | tei63 | 17050.4748 | 20000 | 18525.2374 | 0.0017224 | 9.3388e-05 | 0.047173
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 0 | tei00 | 39.2625 | 43.9398 | 41.6011 | 0.0013412 | 0.00075993 | 0.0046716
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 1 | tei01 | 43.9398 | 48.4219 | 46.1809 | 0.0013412 | 0.00020471 | 0.056465
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 2 | tei02 | 48.4219 | 52.7317 | 50.5768 | 0.0013412 | 0.0001206 | 0.092309
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 3 | tei03 | 52.7317 | 57.9075 | 55.3196 | 0.0013412 | 0.00010884 | 0.13221
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 4 | tei04 | 57.9075 | 64.8216 | 61.3645 | 0.0013412 | 0.00010295 | 0.13125
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 5 | tei05 | 64.8216 | 70.4957 | 67.6587 | 0.0013412 | 9.2937e-05 | 0.20824
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 6 | tei06 | 70.4957 | 77.7766 | 74.1361 | 0.0013412 | 8.9755e-05 | 0.22947
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 7 | tei07 | 77.7766 | 85.6695 | 81.723 | 0.0013412 | 9.6177e-05 | 0.24791
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 8 | tei08 | 85.6695 | 94.1679 | 89.9187 | 0.0013412 | 8.7869e-05 | 0.26125
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 9 | tei09 | 94.1679 | 104.0937 | 99.1308 | 0.0013412 | 7.948e-05 | 0.27542
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 10 | tei10 | 104.0937 | 114.5947 | 109.3442 | 0.0013412 | 8.2772e-05 | 0.28416
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 11 | tei11 | 114.5947 | 126.4362 | 120.5154 | 0.0013412 | 8.6811e-05 | 0.28932
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 12 | tei12 | 126.4362 | 139.5053 | 132.9707 | 0.0013412 | 9.057e-05 | 0.27851
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 13 | tei13 | 139.5053 | 153.6403 | 146.5728 | 0.0013412 | 8.6626e-05 | 0.2691
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 14 | tei14 | 153.6403 | 169.3773 | 161.5088 | 0.0013412 | 8.3615e-05 | 0.25183
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 15 | tei15 | 169.3773 | 186.5858 | 177.9816 | 0.0013412 | 7.9228e-05 | 0.23974
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 16 | tei16 | 186.5858 | 205.4507 | 196.0183 | 0.0013412 | 8.0592e-05 | 0.22846
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 17 | tei17 | 205.4507 | 227.054 | 216.2524 | 0.0013412 | 7.6556e-05 | 0.21408
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 18 | tei18 | 227.054 | 250.9369 | 238.9955 | 0.0013412 | 7.2677e-05 | 0.20386
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 19 | tei19 | 250.9369 | 275.8583 | 263.3976 | 0.0013412 | 6.1126e-05 | 0.19442
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 20 | tei20 | 275.8583 | 303.6288 | 289.7436 | 0.0013412 | 7.4597e-05 | 0.1825
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 21 | tei21 | 303.6288 | 334.1802 | 318.9045 | 0.0013412 | 7.4597e-05 | 0.1712
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 22 | tei22 | 334.1802 | 368.275 | 351.2276 | 0.0013412 | 7.4597e-05 | 0.15765
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 23 | tei23 | 368.275 | 405.8295 | 387.0522 | 0.0013412 | 7.4597e-05 | 0.14756
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 24 | tei24 | 405.8295 | 447.2136 | 426.5215 | 0.0013412 | 7.4597e-05 | 0.13545
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 25 | tei25 | 447.2136 | 492.8178 | 470.0157 | 0.0013412 | 7.4597e-05 | 0.1265
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 26 | tei26 | 492.8178 | 543.0725 | 517.9452 | 0.0013412 | 7.4597e-05 | 0.11568
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 27 | tei27 | 543.0725 | 598.4519 | 570.7622 | 0.0013412 | 7.4597e-05 | 0.1059
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 28 | tei28 | 598.4519 | 659.4785 | 628.9652 | 0.0013412 | 7.4597e-05 | 0.09436
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 29 | tei29 | 659.4785 | 726.7282 | 693.1033 | 0.0013412 | 7.4597e-05 | 0.087687
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 30 | tei30 | 726.7282 | 800.8357 | 763.782 | 0.0013412 | 7.4597e-05 | 0.079259
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 31 | tei31 | 800.8357 | 882.5003 | 841.668 | 0.0013412 | 7.4597e-05 | 0.073274
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 32 | tei32 | 882.5003 | 972.4925 | 927.4964 | 0.0013412 | 7.4597e-05 | 0.065584
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 33 | tei33 | 972.4925 | 1071.6616 | 1022.077 | 0.0013412 | 7.4597e-05 | 0.058248
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 34 | tei34 | 1071.6616 | 1180.9433 | 1126.3025 | 0.0013412 | 7.4597e-05 | 0.051242
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 35 | tei35 | 1180.9433 | 1301.369 | 1241.1562 | 0.0013412 | 7.4597e-05 | 0.044545
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 36 | tei36 | 1301.369 | 1444.1932 | 1372.7811 | 0.0013412 | 7.4597e-05 | 0.03992
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 37 | tei37 | 1444.1932 | 1599.8498 | 1522.0215 | 0.0013412 | 7.4597e-05 | 0.037253
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 38 | tei38 | 1599.8498 | 1763.0124 | 1681.4311 | 0.0013412 | 7.4597e-05 | 0.032993
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 39 | tei39 | 1763.0124 | 1937.1886 | 1850.1005 | 0.0013412 | 7.4597e-05 | 0.028933
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 40 | tei40 | 1937.1886 | 2125.886 | 2031.5373 | 0.0013412 | 7.4597e-05 | 0.025042
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 41 | tei41 | 2125.886 | 2332.6122 | 2229.2491 | 0.0013412 | 7.4597e-05 | 0.021299
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 42 | tei42 | 2332.6122 | 2560.8748 | 2446.7435 | 0.0013412 | 7.4597e-05 | 0.019301
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 43 | tei43 | 2560.8748 | 2814.1814 | 2687.5281 | 0.0013412 | 7.4597e-05 | 0.017372
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 44 | tei44 | 2814.1814 | 3096.0396 | 2955.1105 | 0.0013412 | 7.4597e-05 | 0.015509
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 45 | tei45 | 3096.0396 | 3409.957 | 3252.9983 | 0.0013412 | 7.4597e-05 | 0.013708
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 46 | tei46 | 3409.957 | 3759.4412 | 3584.6991 | 0.0013412 | 7.4597e-05 | 0.011969
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 47 | tei47 | 3759.4412 | 4147.9998 | 3953.7205 | 0.0013412 | 7.4597e-05 | 0.010288
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 48 | tei48 | 4147.9998 | 4579.1404 | 4363.5701 | 0.0013412 | 7.4597e-05 | 0.0086651
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 49 | tei49 | 4579.1404 | 5056.3706 | 4817.7555 | 0.0013412 | 7.4597e-05 | 0.0073807
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 50 | tei50 | 5056.3706 | 5583.198 | 5319.7843 | 0.0013412 | 7.4597e-05 | 0.0062792
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 51 | tei51 | 5583.198 | 6163.1302 | 5873.1641 | 0.0013412 | 7.4597e-05 | 0.0057633
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 52 | tei52 | 6163.1302 | 6799.6748 | 6481.4025 | 0.0013412 | 7.4597e-05 | 0.0049947
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 53 | tei53 | 6799.6748 | 7496.3394 | 7148.0071 | 0.0013412 | 7.4597e-05 | 0.0042511
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 54 | tei54 | 7496.3394 | 8256.6316 | 7876.4855 | 0.0013412 | 7.4597e-05 | 0.0037928
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 55 | tei55 | 8256.6316 | 9084.059 | 8670.3453 | 0.0013412 | 7.4597e-05 | 0.0030914
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 56 | tei56 | 9084.059 | 9982.1292 | 9533.0941 | 0.0013412 | 7.4597e-05 | 0.0027921
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 57 | tei57 | 9982.1292 | 10954.3498 | 10468.2395 | 0.0013412 | 7.4597e-05 | 0.0023768
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 58 | tei58 | 10954.3498 | 12004.2284 | 11479.2891 | 0.0013412 | 7.4597e-05 | 0.0020971
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 59 | tei59 | 12004.2284 | 13135.2726 | 12569.7505 | 0.0013412 | 7.4597e-05 | 0.0017037
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 60 | tei60 | 13135.2726 | 14350.99 | 13743.1313 | 0.0013412 | 7.4597e-05 | 0.066054
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 61 | tei61 | 14350.99 | 15654.8882 | 15002.9391 | 0.0013412 | 7.4597e-05 | 0.060472
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 62 | tei62 | 15654.8882 | 17050.4748 | 16352.6815 | 0.0013412 | 7.4597e-05 | 0.053871
TOFxE Ion | TOFxE_Ion | 414734467.1853 | 662731267.1829 | Ions | T5_Counts | 63 | tei63 | 17050.4748 | 20000 | 18525.2374 | 0.0013412 | 7.4597e-05 | 0.047173

## TOFxE_nonH-POST He extension
Product Description | Product Name | StartET | StopET | Species | Field Name | Channel | ChName | E_Low (Kev) | E_Hi (Kev) | Mid_E(KeV) | G (cm^2 sr) | G-small | efficiency
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 0 | ten00 | 56.8491 | 73.6576 | 65 | 0.0017196 | 0.00024285 | 0.29
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 1 | ten01 | 73.6576 | 95.4358 | 85 | 0.0017196 | 0.00012142 | 0.72
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 2 | ten02 | 95.4358 | 123.6532 | 110 | 0.0017196 | 0.00012142 | 0.83
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 3 | ten03 | 123.6532 | 160.2135 | 142 | 0.0017196 | 0.00012142 | 0.9
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 4 | ten04 | 160.2135 | 207.5836 | 184 | 0.0017196 | 0.00012142 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 5 | ten05 | 207.5836 | 268.9595 | 238 | 0.0017196 | 0.00012142 | 0.96
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 6 | ten06 | 268.9595 | 348.4824 | 309 | 0.0017196 | 0.00012142 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 7 | ten07 | 348.4824 | 451.5176 | 400 | 0.0017196 | 0.00012142 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T0_Counts | 8 | ten08 | 451.5176 | 585.0171 | 518 | 0.0017196 | 0.00012142 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T0_Counts | 9 | ten09 | 585.0171 | 757.9882 | 671 | 0.0017196 | 0.00012142 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T0_Counts | 10 | ten10 | 757.9882 | 982.0118 | 870 | 0.0017196 | 0.00012142 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T0_Counts | 11 | ten11 | 123.802 | 160.3236 | 142 | 0.0017196 | 0.00012142 | 0.3
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T0_Counts | 12 | ten12 | 160.3236 | 207.619 | 184 | 0.0017196 | 0.00012142 | 0.7
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T0_Counts | 13 | ten13 | 207.619 | 268.8667 | 238 | 0.0017196 | 0.00012142 | 0.85
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T0_Counts | 14 | ten14 | 268.8667 | 348.1823 | 309 | 0.0017196 | 0.00012142 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T0_Counts | 15 | ten15 | 348.1823 | 450.8961 | 400 | 0.0017196 | 0.00012142 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T0_Counts | 16 | ten16 | 450.8961 | 583.9105 | 518 | 0.0017196 | 0.00012142 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T0_Counts | 17 | ten17 | 594.5 | 770.5 | 671 | 0.0017196 | 0.00012142 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T0_Counts | 18 | ten18 | 770.5 | 998.5 | 870 | 0.0017196 | 0.00012142 | 1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 662731267.1829 | X | T0_Counts | 19 | ten19 | -1 | -1 | -1 | 0.0017196 | 0.00012142 | -1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 0 | ten00 | 56.8491 | 73.6576 | 65 | 0.0017478 | 0.00023975 | 0.29
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 1 | ten01 | 73.6576 | 95.4358 | 85 | 0.0017478 | 0.00011988 | 0.72
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 2 | ten02 | 95.4358 | 123.6532 | 110 | 0.0017478 | 0.00011988 | 0.83
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 3 | ten03 | 123.6532 | 160.2135 | 142 | 0.0017478 | 0.00011988 | 0.9
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 4 | ten04 | 160.2135 | 207.5836 | 184 | 0.0017478 | 0.00011988 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 5 | ten05 | 207.5836 | 268.9595 | 238 | 0.0017478 | 0.00011988 | 0.96
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 6 | ten06 | 268.9595 | 348.4824 | 309 | 0.0017478 | 0.00011988 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 7 | ten07 | 348.4824 | 451.5176 | 400 | 0.0017478 | 0.00011988 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T1_Counts | 8 | ten08 | 451.5176 | 585.0171 | 518 | 0.0017478 | 0.00011988 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T1_Counts | 9 | ten09 | 585.0171 | 757.9882 | 671 | 0.0017478 | 0.00011988 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T1_Counts | 10 | ten10 | 757.9882 | 982.0118 | 870 | 0.0017478 | 0.00011988 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T1_Counts | 11 | ten11 | 123.802 | 160.3236 | 142 | 0.0017478 | 0.00011988 | 0.3
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T1_Counts | 12 | ten12 | 160.3236 | 207.619 | 184 | 0.0017478 | 0.00011988 | 0.7
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T1_Counts | 13 | ten13 | 207.619 | 268.8667 | 238 | 0.0017478 | 0.00011988 | 0.85
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T1_Counts | 14 | ten14 | 268.8667 | 348.1823 | 309 | 0.0017478 | 0.00011988 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T1_Counts | 15 | ten15 | 348.1823 | 450.8961 | 400 | 0.0017478 | 0.00011988 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T1_Counts | 16 | ten16 | 450.8961 | 583.9105 | 518 | 0.0017478 | 0.00011988 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T1_Counts | 17 | ten17 | 594.5 | 770.5 | 671 | 0.0017478 | 0.00011988 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T1_Counts | 18 | ten18 | 770.5 | 998.5 | 870 | 0.0017478 | 0.00011988 | 1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 662731267.1829 | X | T1_Counts | 19 | ten19 | -1 | -1 | -1 | 0.0017478 | 0.00011988 | -1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 0 | ten00 | 56.8491 | 73.6576 | 65 | 0.0019402 | 0.00026084 | 0.29
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 1 | ten01 | 73.6576 | 95.4358 | 85 | 0.0019402 | 0.00013042 | 0.72
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 2 | ten02 | 95.4358 | 123.6532 | 110 | 0.0019402 | 0.00013042 | 0.83
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 3 | ten03 | 123.6532 | 160.2135 | 142 | 0.0019402 | 0.00013042 | 0.9
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 4 | ten04 | 160.2135 | 207.5836 | 184 | 0.0019402 | 0.00013042 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 5 | ten05 | 207.5836 | 268.9595 | 238 | 0.0019402 | 0.00013042 | 0.96
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 6 | ten06 | 268.9595 | 348.4824 | 309 | 0.0019402 | 0.00013042 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 7 | ten07 | 348.4824 | 451.5176 | 400 | 0.0019402 | 0.00013042 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T2_Counts | 8 | ten08 | 451.5176 | 585.0171 | 518 | 0.0019402 | 0.00013042 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T2_Counts | 9 | ten09 | 585.0171 | 757.9882 | 671 | 0.0019402 | 0.00013042 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T2_Counts | 10 | ten10 | 757.9882 | 982.0118 | 870 | 0.0019402 | 0.00013042 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T2_Counts | 11 | ten11 | 123.802 | 160.3236 | 142 | 0.0019402 | 0.00013042 | 0.3
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T2_Counts | 12 | ten12 | 160.3236 | 207.619 | 184 | 0.0019402 | 0.00013042 | 0.7
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T2_Counts | 13 | ten13 | 207.619 | 268.8667 | 238 | 0.0019402 | 0.00013042 | 0.85
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T2_Counts | 14 | ten14 | 268.8667 | 348.1823 | 309 | 0.0019402 | 0.00013042 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T2_Counts | 15 | ten15 | 348.1823 | 450.8961 | 400 | 0.0019402 | 0.00013042 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T2_Counts | 16 | ten16 | 450.8961 | 583.9105 | 518 | 0.0019402 | 0.00013042 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T2_Counts | 17 | ten17 | 594.5 | 770.5 | 671 | 0.0019402 | 0.00013042 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T2_Counts | 18 | ten18 | 770.5 | 998.5 | 870 | 0.0019402 | 0.00013042 | 1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 662731267.1829 | X | T2_Counts | 19 | ten19 | -1 | -1 | -1 | 0.0019402 | 0.00013042 | -1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 0 | ten00 | 56.8491 | 73.6576 | 65 | 0.001252 | 0.00017088 | 0.29
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 1 | ten01 | 73.6576 | 95.4358 | 85 | 0.001252 | 8.5442e-05 | 0.72
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 2 | ten02 | 95.4358 | 123.6532 | 110 | 0.001252 | 8.5442e-05 | 0.83
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 3 | ten03 | 123.6532 | 160.2135 | 142 | 0.001252 | 8.5442e-05 | 0.9
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 4 | ten04 | 160.2135 | 207.5836 | 184 | 0.001252 | 8.5442e-05 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 5 | ten05 | 207.5836 | 268.9595 | 238 | 0.001252 | 8.5442e-05 | 0.96
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 6 | ten06 | 268.9595 | 348.4824 | 309 | 0.001252 | 8.5442e-05 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 7 | ten07 | 348.4824 | 451.5176 | 400 | 0.001252 | 8.5442e-05 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T3_Counts | 8 | ten08 | 451.5176 | 585.0171 | 518 | 0.001252 | 8.5442e-05 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T3_Counts | 9 | ten09 | 585.0171 | 757.9882 | 671 | 0.001252 | 8.5442e-05 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T3_Counts | 10 | ten10 | 757.9882 | 982.0118 | 870 | 0.001252 | 8.5442e-05 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T3_Counts | 11 | ten11 | 123.802 | 160.3236 | 142 | 0.001252 | 8.5442e-05 | 0.3
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T3_Counts | 12 | ten12 | 160.3236 | 207.619 | 184 | 0.001252 | 8.5442e-05 | 0.7
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T3_Counts | 13 | ten13 | 207.619 | 268.8667 | 238 | 0.001252 | 8.5442e-05 | 0.85
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T3_Counts | 14 | ten14 | 268.8667 | 348.1823 | 309 | 0.001252 | 8.5442e-05 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T3_Counts | 15 | ten15 | 348.1823 | 450.8961 | 400 | 0.001252 | 8.5442e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T3_Counts | 16 | ten16 | 450.8961 | 583.9105 | 518 | 0.001252 | 8.5442e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T3_Counts | 17 | ten17 | 594.5 | 770.5 | 671 | 0.001252 | 8.5442e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T3_Counts | 18 | ten18 | 770.5 | 998.5 | 870 | 0.001252 | 8.5442e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 662731267.1829 | X | T3_Counts | 19 | ten19 | -1 | -1 | -1 | 0.001252 | 8.5442e-05 | -1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 0 | ten00 | 56.8491 | 73.6576 | 65 | 0.0017224 | 0.00022967 | 0.29
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 1 | ten01 | 73.6576 | 95.4358 | 85 | 0.0017224 | 0.00011483 | 0.72
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 2 | ten02 | 95.4358 | 123.6532 | 110 | 0.0017224 | 0.00011483 | 0.83
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 3 | ten03 | 123.6532 | 160.2135 | 142 | 0.0017224 | 0.00011483 | 0.9
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 4 | ten04 | 160.2135 | 207.5836 | 184 | 0.0017224 | 0.00011483 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 5 | ten05 | 207.5836 | 268.9595 | 238 | 0.0017224 | 0.00011483 | 0.96
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 6 | ten06 | 268.9595 | 348.4824 | 309 | 0.0017224 | 0.00011483 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 7 | ten07 | 348.4824 | 451.5176 | 400 | 0.0017224 | 0.00011483 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T4_Counts | 8 | ten08 | 451.5176 | 585.0171 | 518 | 0.0017224 | 0.00011483 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T4_Counts | 9 | ten09 | 585.0171 | 757.9882 | 671 | 0.0017224 | 0.00011483 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T4_Counts | 10 | ten10 | 757.9882 | 982.0118 | 870 | 0.0017224 | 0.00011483 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T4_Counts | 11 | ten11 | 123.802 | 160.3236 | 142 | 0.0017224 | 0.00011483 | 0.3
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T4_Counts | 12 | ten12 | 160.3236 | 207.619 | 184 | 0.0017224 | 0.00011483 | 0.7
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T4_Counts | 13 | ten13 | 207.619 | 268.8667 | 238 | 0.0017224 | 0.00011483 | 0.85
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T4_Counts | 14 | ten14 | 268.8667 | 348.1823 | 309 | 0.0017224 | 0.00011483 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T4_Counts | 15 | ten15 | 348.1823 | 450.8961 | 400 | 0.0017224 | 0.00011483 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T4_Counts | 16 | ten16 | 450.8961 | 583.9105 | 518 | 0.0017224 | 0.00011483 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T4_Counts | 17 | ten17 | 594.5 | 770.5 | 671 | 0.0017224 | 0.00011483 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T4_Counts | 18 | ten18 | 770.5 | 998.5 | 870 | 0.0017224 | 0.00011483 | 1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 662731267.1829 | X | T4_Counts | 19 | ten19 | -1 | -1 | -1 | 0.0017224 | 0.00011483 | -1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 0 | ten00 | 56.8491 | 73.6576 | 65 | 0.0013412 | 0.00018346 | 0.29
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 1 | ten01 | 73.6576 | 95.4358 | 85 | 0.0013412 | 9.1728e-05 | 0.72
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 2 | ten02 | 95.4358 | 123.6532 | 110 | 0.0013412 | 9.1728e-05 | 0.83
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 3 | ten03 | 123.6532 | 160.2135 | 142 | 0.0013412 | 9.1728e-05 | 0.9
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 4 | ten04 | 160.2135 | 207.5836 | 184 | 0.0013412 | 9.1728e-05 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 5 | ten05 | 207.5836 | 268.9595 | 238 | 0.0013412 | 9.1728e-05 | 0.96
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 6 | ten06 | 268.9595 | 348.4824 | 309 | 0.0013412 | 9.1728e-05 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 7 | ten07 | 348.4824 | 451.5176 | 400 | 0.0013412 | 9.1728e-05 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T5_Counts | 8 | ten08 | 451.5176 | 585.0171 | 518 | 0.0013412 | 9.1728e-05 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T5_Counts | 9 | ten09 | 585.0171 | 757.9882 | 671 | 0.0013412 | 9.1728e-05 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | He | T5_Counts | 10 | ten10 | 757.9882 | 982.0118 | 870 | 0.0013412 | 9.1728e-05 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T5_Counts | 11 | ten11 | 123.802 | 160.3236 | 142 | 0.0013412 | 9.1728e-05 | 0.3
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T5_Counts | 12 | ten12 | 160.3236 | 207.619 | 184 | 0.0013412 | 9.1728e-05 | 0.7
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T5_Counts | 13 | ten13 | 207.619 | 268.8667 | 238 | 0.0013412 | 9.1728e-05 | 0.85
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T5_Counts | 14 | ten14 | 268.8667 | 348.1823 | 309 | 0.0013412 | 9.1728e-05 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T5_Counts | 15 | ten15 | 348.1823 | 450.8961 | 400 | 0.0013412 | 9.1728e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T5_Counts | 16 | ten16 | 450.8961 | 583.9105 | 518 | 0.0013412 | 9.1728e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T5_Counts | 17 | ten17 | 594.5 | 770.5 | 671 | 0.0013412 | 9.1728e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 474523350.1843 | 662731267.1829 | O | T5_Counts | 18 | ten18 | 770.5 | 998.5 | 870 | 0.0013412 | 9.1728e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 662731267.1829 | X | T5_Counts | 19 | ten19 | -1 | -1 | -1 | 0.0013412 | 9.1728e-05 | -1

## TOFxE_nonH-PRE He extension
Product Description | Product Name | StartET | StopET | Species | Field Name | Channel | ChName | E_Low (Kev) | E_Hi (Kev) | Mid_E(KeV) | G (cm^2 sr) | G-small | efficiency
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 0 | ten00 | 56.8491 | 73.6576 | 65 | 0.0017196 | 0.00024285 | 0.29
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 1 | ten01 | 73.6576 | 95.4358 | 85 | 0.0017196 | 0.00012142 | 0.72
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 2 | ten02 | 95.4358 | 123.6532 | 110 | 0.0017196 | 0.00012142 | 0.83
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 3 | ten03 | 123.6532 | 160.2135 | 142 | 0.0017196 | 0.00012142 | 0.9
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 4 | ten04 | 160.2135 | 207.5836 | 184 | 0.0017196 | 0.00012142 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 5 | ten05 | 207.5836 | 268.9595 | 238 | 0.0017196 | 0.00012142 | 0.96
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 6 | ten06 | 268.9595 | 348.4824 | 309 | 0.0017196 | 0.00012142 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T0_Counts | 7 | ten07 | 348.4824 | 451.5176 | 400 | 0.0017196 | 0.00012142 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | He | T0_Counts | 8 | ten08 | 451.5176 | 584.4824 | 518 | 0.0017196 | 0.00012142 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T0_Counts | 9 | ten09 | 123.802 | 160.3236 | 142 | 0.0017196 | 0.00012142 | 0.3
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T0_Counts | 10 | ten10 | 160.3236 | 207.619 | 184 | 0.0017196 | 0.00012142 | 0.7
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T0_Counts | 11 | ten11 | 207.619 | 268.8667 | 238 | 0.0017196 | 0.00012142 | 0.85
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T0_Counts | 12 | ten12 | 268.8667 | 348.1823 | 309 | 0.0017196 | 0.00012142 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T0_Counts | 13 | ten13 | 348.1823 | 450.8961 | 400 | 0.0017196 | 0.00012142 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T0_Counts | 14 | ten14 | 450.8961 | 583.9105 | 518 | 0.0017196 | 0.00012142 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T0_Counts | 15 | ten15 | 594.5 | 770.5 | 671 | 0.0017196 | 0.00012142 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T0_Counts | 16 | ten16 | 770.5 | 998.5 | 870 | 0.0017196 | 0.00012142 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T0_Counts | 17 | ten17 | 998.5 | 1256 | 1127 | 0.0017196 | 0.00012142 | 1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 474523350.1843 | X | T0_Counts | 18 | ten18 | -1 | -1 | -1 | 0.0017196 | 0.00012142 | -1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 662731267.1829 | X | T0_Counts | 19 | ten19 | -1 | -1 | -1 | 0.0017196 | 0.00012142 | -1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 0 | ten00 | 56.8491 | 73.6576 | 65 | 0.0017478 | 0.00023975 | 0.29
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 1 | ten01 | 73.6576 | 95.4358 | 85 | 0.0017478 | 0.00011988 | 0.72
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 2 | ten02 | 95.4358 | 123.6532 | 110 | 0.0017478 | 0.00011988 | 0.83
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 3 | ten03 | 123.6532 | 160.2135 | 142 | 0.0017478 | 0.00011988 | 0.9
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 4 | ten04 | 160.2135 | 207.5836 | 184 | 0.0017478 | 0.00011988 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 5 | ten05 | 207.5836 | 268.9595 | 238 | 0.0017478 | 0.00011988 | 0.96
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 6 | ten06 | 268.9595 | 348.4824 | 309 | 0.0017478 | 0.00011988 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T1_Counts | 7 | ten07 | 348.4824 | 451.5176 | 400 | 0.0017478 | 0.00011988 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | He | T1_Counts | 8 | ten08 | 451.5176 | 584.4824 | 518 | 0.0017478 | 0.00011988 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T1_Counts | 9 | ten09 | 123.802 | 160.3236 | 142 | 0.0017478 | 0.00011988 | 0.3
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T1_Counts | 10 | ten10 | 160.3236 | 207.619 | 184 | 0.0017478 | 0.00011988 | 0.7
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T1_Counts | 11 | ten11 | 207.619 | 268.8667 | 238 | 0.0017478 | 0.00011988 | 0.85
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T1_Counts | 12 | ten12 | 268.8667 | 348.1823 | 309 | 0.0017478 | 0.00011988 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T1_Counts | 13 | ten13 | 348.1823 | 450.8961 | 400 | 0.0017478 | 0.00011988 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T1_Counts | 14 | ten14 | 450.8961 | 583.9105 | 518 | 0.0017478 | 0.00011988 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T1_Counts | 15 | ten15 | 594.5 | 770.5 | 671 | 0.0017478 | 0.00011988 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T1_Counts | 16 | ten16 | 770.5 | 998.5 | 870 | 0.0017478 | 0.00011988 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T1_Counts | 17 | ten17 | 998.5 | 1256 | 1127 | 0.0017478 | 0.00011988 | 1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 474523350.1843 | X | T1_Counts | 18 | ten18 | -1 | -1 | -1 | 0.0017478 | 0.00011988 | -1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 662731267.1829 | X | T1_Counts | 19 | ten19 | -1 | -1 | -1 | 0.0017478 | 0.00011988 | -1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 0 | ten00 | 56.8491 | 73.6576 | 65 | 0.0019402 | 0.00026084 | 0.29
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 1 | ten01 | 73.6576 | 95.4358 | 85 | 0.0019402 | 0.00013042 | 0.72
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 2 | ten02 | 95.4358 | 123.6532 | 110 | 0.0019402 | 0.00013042 | 0.83
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 3 | ten03 | 123.6532 | 160.2135 | 142 | 0.0019402 | 0.00013042 | 0.9
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 4 | ten04 | 160.2135 | 207.5836 | 184 | 0.0019402 | 0.00013042 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 5 | ten05 | 207.5836 | 268.9595 | 238 | 0.0019402 | 0.00013042 | 0.96
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 6 | ten06 | 268.9595 | 348.4824 | 309 | 0.0019402 | 0.00013042 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T2_Counts | 7 | ten07 | 348.4824 | 451.5176 | 400 | 0.0019402 | 0.00013042 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | He | T2_Counts | 8 | ten08 | 451.5176 | 584.4824 | 518 | 0.0019402 | 0.00013042 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T2_Counts | 9 | ten09 | 123.802 | 160.3236 | 142 | 0.0019402 | 0.00013042 | 0.3
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T2_Counts | 10 | ten10 | 160.3236 | 207.619 | 184 | 0.0019402 | 0.00013042 | 0.7
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T2_Counts | 11 | ten11 | 207.619 | 268.8667 | 238 | 0.0019402 | 0.00013042 | 0.85
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T2_Counts | 12 | ten12 | 268.8667 | 348.1823 | 309 | 0.0019402 | 0.00013042 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T2_Counts | 13 | ten13 | 348.1823 | 450.8961 | 400 | 0.0019402 | 0.00013042 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T2_Counts | 14 | ten14 | 450.8961 | 583.9105 | 518 | 0.0019402 | 0.00013042 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T2_Counts | 15 | ten15 | 594.5 | 770.5 | 671 | 0.0019402 | 0.00013042 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T2_Counts | 16 | ten16 | 770.5 | 998.5 | 870 | 0.0019402 | 0.00013042 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T2_Counts | 17 | ten17 | 998.5 | 1256 | 1127 | 0.0019402 | 0.00013042 | 1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 474523350.1843 | X | T2_Counts | 18 | ten18 | -1 | -1 | -1 | 0.0019402 | 0.00013042 | -1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 662731267.1829 | X | T2_Counts | 19 | ten19 | -1 | -1 | -1 | 0.0019402 | 0.00013042 | -1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 0 | ten00 | 56.8491 | 73.6576 | 65 | 0.001252 | 0.00017088 | 0.29
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 1 | ten01 | 73.6576 | 95.4358 | 85 | 0.001252 | 8.5442e-05 | 0.72
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 2 | ten02 | 95.4358 | 123.6532 | 110 | 0.001252 | 8.5442e-05 | 0.83
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 3 | ten03 | 123.6532 | 160.2135 | 142 | 0.001252 | 8.5442e-05 | 0.9
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 4 | ten04 | 160.2135 | 207.5836 | 184 | 0.001252 | 8.5442e-05 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 5 | ten05 | 207.5836 | 268.9595 | 238 | 0.001252 | 8.5442e-05 | 0.96
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 6 | ten06 | 268.9595 | 348.4824 | 309 | 0.001252 | 8.5442e-05 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T3_Counts | 7 | ten07 | 348.4824 | 451.5176 | 400 | 0.001252 | 8.5442e-05 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | He | T3_Counts | 8 | ten08 | 451.5176 | 584.4824 | 518 | 0.001252 | 8.5442e-05 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T3_Counts | 9 | ten09 | 123.802 | 160.3236 | 142 | 0.001252 | 8.5442e-05 | 0.3
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T3_Counts | 10 | ten10 | 160.3236 | 207.619 | 184 | 0.001252 | 8.5442e-05 | 0.7
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T3_Counts | 11 | ten11 | 207.619 | 268.8667 | 238 | 0.001252 | 8.5442e-05 | 0.85
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T3_Counts | 12 | ten12 | 268.8667 | 348.1823 | 309 | 0.001252 | 8.5442e-05 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T3_Counts | 13 | ten13 | 348.1823 | 450.8961 | 400 | 0.001252 | 8.5442e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T3_Counts | 14 | ten14 | 450.8961 | 583.9105 | 518 | 0.001252 | 8.5442e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T3_Counts | 15 | ten15 | 594.5 | 770.5 | 671 | 0.001252 | 8.5442e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T3_Counts | 16 | ten16 | 770.5 | 998.5 | 870 | 0.001252 | 8.5442e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T3_Counts | 17 | ten17 | 998.5 | 1256 | 1127 | 0.001252 | 8.5442e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 474523350.1843 | X | T3_Counts | 18 | ten18 | -1 | -1 | -1 | 0.001252 | 8.5442e-05 | -1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 662731267.1829 | X | T3_Counts | 19 | ten19 | -1 | -1 | -1 | 0.001252 | 8.5442e-05 | -1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 0 | ten00 | 56.8491 | 73.6576 | 65 | 0.0017224 | 0.00022967 | 0.29
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 1 | ten01 | 73.6576 | 95.4358 | 85 | 0.0017224 | 0.00011483 | 0.72
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 2 | ten02 | 95.4358 | 123.6532 | 110 | 0.0017224 | 0.00011483 | 0.83
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 3 | ten03 | 123.6532 | 160.2135 | 142 | 0.0017224 | 0.00011483 | 0.9
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 4 | ten04 | 160.2135 | 207.5836 | 184 | 0.0017224 | 0.00011483 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 5 | ten05 | 207.5836 | 268.9595 | 238 | 0.0017224 | 0.00011483 | 0.96
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 6 | ten06 | 268.9595 | 348.4824 | 309 | 0.0017224 | 0.00011483 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T4_Counts | 7 | ten07 | 348.4824 | 451.5176 | 400 | 0.0017224 | 0.00011483 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | He | T4_Counts | 8 | ten08 | 451.5176 | 584.4824 | 518 | 0.0017224 | 0.00011483 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T4_Counts | 9 | ten09 | 123.802 | 160.3236 | 142 | 0.0017224 | 0.00011483 | 0.3
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T4_Counts | 10 | ten10 | 160.3236 | 207.619 | 184 | 0.0017224 | 0.00011483 | 0.7
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T4_Counts | 11 | ten11 | 207.619 | 268.8667 | 238 | 0.0017224 | 0.00011483 | 0.85
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T4_Counts | 12 | ten12 | 268.8667 | 348.1823 | 309 | 0.0017224 | 0.00011483 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T4_Counts | 13 | ten13 | 348.1823 | 450.8961 | 400 | 0.0017224 | 0.00011483 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T4_Counts | 14 | ten14 | 450.8961 | 583.9105 | 518 | 0.0017224 | 0.00011483 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T4_Counts | 15 | ten15 | 594.5 | 770.5 | 671 | 0.0017224 | 0.00011483 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T4_Counts | 16 | ten16 | 770.5 | 998.5 | 870 | 0.0017224 | 0.00011483 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T4_Counts | 17 | ten17 | 998.5 | 1256 | 1127 | 0.0017224 | 0.00011483 | 1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 474523350.1843 | X | T4_Counts | 18 | ten18 | -1 | -1 | -1 | 0.0017224 | 0.00011483 | -1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 662731267.1829 | X | T4_Counts | 19 | ten19 | -1 | -1 | -1 | 0.0017224 | 0.00011483 | -1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 0 | ten00 | 56.8491 | 73.6576 | 65 | 0.0013412 | 0.00018346 | 0.29
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 1 | ten01 | 73.6576 | 95.4358 | 85 | 0.0013412 | 9.1728e-05 | 0.72
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 2 | ten02 | 95.4358 | 123.6532 | 110 | 0.0013412 | 9.1728e-05 | 0.83
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 3 | ten03 | 123.6532 | 160.2135 | 142 | 0.0013412 | 9.1728e-05 | 0.9
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 4 | ten04 | 160.2135 | 207.5836 | 184 | 0.0013412 | 9.1728e-05 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 5 | ten05 | 207.5836 | 268.9595 | 238 | 0.0013412 | 9.1728e-05 | 0.96
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 6 | ten06 | 268.9595 | 348.4824 | 309 | 0.0013412 | 9.1728e-05 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 662731267.1829 | He | T5_Counts | 7 | ten07 | 348.4824 | 451.5176 | 400 | 0.0013412 | 9.1728e-05 | 0.98
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | He | T5_Counts | 8 | ten08 | 451.5176 | 584.4824 | 518 | 0.0013412 | 9.1728e-05 | 0.97
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T5_Counts | 9 | ten09 | 123.802 | 160.3236 | 142 | 0.0013412 | 9.1728e-05 | 0.3
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T5_Counts | 10 | ten10 | 160.3236 | 207.619 | 184 | 0.0013412 | 9.1728e-05 | 0.7
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T5_Counts | 11 | ten11 | 207.619 | 268.8667 | 238 | 0.0013412 | 9.1728e-05 | 0.85
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T5_Counts | 12 | ten12 | 268.8667 | 348.1823 | 309 | 0.0013412 | 9.1728e-05 | 0.93
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T5_Counts | 13 | ten13 | 348.1823 | 450.8961 | 400 | 0.0013412 | 9.1728e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T5_Counts | 14 | ten14 | 450.8961 | 583.9105 | 518 | 0.0013412 | 9.1728e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T5_Counts | 15 | ten15 | 594.5 | 770.5 | 671 | 0.0013412 | 9.1728e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T5_Counts | 16 | ten16 | 770.5 | 998.5 | 870 | 0.0013412 | 9.1728e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 414734467.1853 | 474523350.1843 | O | T5_Counts | 17 | ten17 | 998.5 | 1256 | 1127 | 0.0013412 | 9.1728e-05 | 1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 474523350.1843 | X | T5_Counts | 18 | ten18 | -1 | -1 | -1 | 0.0013412 | 9.1728e-05 | -1
TOFxE non Hydrogen | TOFxE_nonH | 347112066.1839 | 662731267.1829 | X | T5_Counts | 19 | ten19 | -1 | -1 | -1 | 0.0013412 | 9.1728e-05 | -1

## TOFxPH_H_LEHT
Product Description | Product Name | StartET | StopET | Species | Field Name | Channel | ChName | E_Low (Kev) | E_Hi (Kev) | Mid_E(KeV) | G (cm^2 sr) | G-small | efficiency
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 0 | tpp00 | 49 | 76 | 63.9414 | 0.0018344 | 0.0018344 | 0.059211
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 1 | tpp01 | 89 | 120 | 105.6433 | 0.0018344 | 0.0018344 | 0.10417
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 2 | tpp02 | 132 | 161 | 147.2158 | 0.0018344 | 0.0018344 | 0.2795
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T0_Counts | 3 | tpp03 | 7 | 9.5 | 8.3442 | 0.0018344 | 0.0018344 | 0.012269
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T0_Counts | 4 | tpp04 | 9.5 | 12.75 | 11.2431 | 0.0018344 | 0.0018344 | 0.030645
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T0_Counts | 5 | tpp05 | 12.75 | 21 | 17.3719 | 0.0018344 | 0.0018344 | 0.05402
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T0_Counts | 6 | tpp06 | 21 | 28 | 24.7487 | 0.0018344 | 0.0018344 | 0.077006
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T0_Counts | 7 | tpp07 | 28 | 34.5 | 31.4185 | 0.0018344 | 0.0018344 | 0.095761
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T0_Counts | 8 | tpp08 | 34.5 | 46.5 | 40.942 | 0.0018344 | 0.0018344 | 0.12462
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T0_Counts | 9 | tpp09 | 47 | 100 | 50 | 0.0018344 | 0.0018344 | 0.255
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 0 | tpp00 | 49 | 76 | 63.9414 | 0.002 | 0.002 | 0.059211
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 1 | tpp01 | 89 | 120 | 105.6433 | 0.002 | 0.002 | 0.10417
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 2 | tpp02 | 132 | 161 | 147.2158 | 0.002 | 0.002 | 0.2795
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T1_Counts | 3 | tpp03 | 7 | 9.5 | 8.3442 | 0.002 | 0.002 | 0.012269
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T1_Counts | 4 | tpp04 | 9.5 | 12.75 | 11.2431 | 0.002 | 0.002 | 0.030645
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T1_Counts | 5 | tpp05 | 12.75 | 21 | 17.3719 | 0.002 | 0.002 | 0.05402
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T1_Counts | 6 | tpp06 | 21 | 28 | 24.7487 | 0.002 | 0.002 | 0.077006
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T1_Counts | 7 | tpp07 | 28 | 34.5 | 31.4185 | 0.002 | 0.002 | 0.095761
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T1_Counts | 8 | tpp08 | 34.5 | 46.5 | 40.942 | 0.002 | 0.002 | 0.12462
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T1_Counts | 9 | tpp09 | 47 | 100 | 50 | 0.002 | 0.002 | 0.255
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 0 | tpp00 | 49 | 76 | 63.9414 | 0.0025083 | 0.0025083 | 0.059211
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 1 | tpp01 | 89 | 120 | 105.6433 | 0.0025083 | 0.0025083 | 0.10417
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 2 | tpp02 | 132 | 161 | 147.2158 | 0.0025083 | 0.0025083 | 0.2795
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T2_Counts | 3 | tpp03 | 7 | 9.5 | 8.3442 | 0.0025083 | 0.0025083 | 0.012269
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T2_Counts | 4 | tpp04 | 9.5 | 12.75 | 11.2431 | 0.0025083 | 0.0025083 | 0.030645
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T2_Counts | 5 | tpp05 | 12.75 | 21 | 17.3719 | 0.0025083 | 0.0025083 | 0.05402
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T2_Counts | 6 | tpp06 | 21 | 28 | 24.7487 | 0.0025083 | 0.0025083 | 0.077006
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T2_Counts | 7 | tpp07 | 28 | 34.5 | 31.4185 | 0.0025083 | 0.0025083 | 0.095761
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T2_Counts | 8 | tpp08 | 34.5 | 46.5 | 40.942 | 0.0025083 | 0.0025083 | 0.12462
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T2_Counts | 9 | tpp09 | 47 | 100 | 50 | 0.0025083 | 0.0025083 | 0.255
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 0 | tpp00 | 49 | 76 | 63.9414 | 0.0013622 | 0.0013622 | 0.059211
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 1 | tpp01 | 89 | 120 | 105.6433 | 0.0013622 | 0.0013622 | 0.10417
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 2 | tpp02 | 132 | 161 | 147.2158 | 0.0013622 | 0.0013622 | 0.2795
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T3_Counts | 3 | tpp03 | 7 | 9.5 | 8.3442 | 0.0013622 | 0.0013622 | 0.012269
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T3_Counts | 4 | tpp04 | 9.5 | 12.75 | 11.2431 | 0.0013622 | 0.0013622 | 0.030645
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T3_Counts | 5 | tpp05 | 12.75 | 21 | 17.3719 | 0.0013622 | 0.0013622 | 0.05402
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T3_Counts | 6 | tpp06 | 21 | 28 | 24.7487 | 0.0013622 | 0.0013622 | 0.077006
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T3_Counts | 7 | tpp07 | 28 | 34.5 | 31.4185 | 0.0013622 | 0.0013622 | 0.095761
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T3_Counts | 8 | tpp08 | 34.5 | 46.5 | 40.942 | 0.0013622 | 0.0013622 | 0.12462
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T3_Counts | 9 | tpp09 | 47 | 100 | 50 | 0.0013622 | 0.0013622 | 0.255
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 0 | tpp00 | 49 | 76 | 63.9414 | 0.0019716 | 0.0019716 | 0.059211
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 1 | tpp01 | 89 | 120 | 105.6433 | 0.0019716 | 0.0019716 | 0.10417
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 2 | tpp02 | 132 | 161 | 147.2158 | 0.0019716 | 0.0019716 | 0.2795
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T4_Counts | 3 | tpp03 | 7 | 9.5 | 8.3442 | 0.0019716 | 0.0019716 | 0.012269
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T4_Counts | 4 | tpp04 | 9.5 | 12.75 | 11.2431 | 0.0019716 | 0.0019716 | 0.030645
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T4_Counts | 5 | tpp05 | 12.75 | 21 | 17.3719 | 0.0019716 | 0.0019716 | 0.05402
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T4_Counts | 6 | tpp06 | 21 | 28 | 24.7487 | 0.0019716 | 0.0019716 | 0.077006
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T4_Counts | 7 | tpp07 | 28 | 34.5 | 31.4185 | 0.0019716 | 0.0019716 | 0.095761
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T4_Counts | 8 | tpp08 | 34.5 | 46.5 | 40.942 | 0.0019716 | 0.0019716 | 0.12462
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T4_Counts | 9 | tpp09 | 47 | 100 | 50 | 0.0019716 | 0.0019716 | 0.255
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 0 | tpp00 | 49 | 76 | 63.9414 | 0.0011646 | 0.0011646 | 0.059211
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 1 | tpp01 | 89 | 120 | 105.6433 | 0.0011646 | 0.0011646 | 0.10417
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 2 | tpp02 | 132 | 161 | 147.2158 | 0.0011646 | 0.0011646 | 0.2795
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T5_Counts | 3 | tpp03 | 7 | 9.5 | 8.3442 | 0.0011646 | 0.0011646 | 0.012269
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T5_Counts | 4 | tpp04 | 9.5 | 12.75 | 11.2431 | 0.0011646 | 0.0011646 | 0.030645
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T5_Counts | 5 | tpp05 | 12.75 | 21 | 17.3719 | 0.0011646 | 0.0011646 | 0.05402
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T5_Counts | 6 | tpp06 | 21 | 28 | 24.7487 | 0.0011646 | 0.0011646 | 0.077006
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T5_Counts | 7 | tpp07 | 28 | 34.5 | 31.4185 | 0.0011646 | 0.0011646 | 0.095761
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T5_Counts | 8 | tpp08 | 34.5 | 46.5 | 40.942 | 0.0011646 | 0.0011646 | 0.12462
TOFxPH Hydrogen - Low Energy Res High Time Res | TOFxPH_H_LEHT | 414734467.1853 | 662731267.1829 | H | T5_Counts | 9 | tpp09 | 47 | 100 | 50 | 0.0011646 | 0.0011646 | 0.255

## TOFxPH_H_HELT
Product Description | Product Name | StartET | StopET | Species | Field Name | Channel | ChName | E_Low (Kev) | E_Hi (Kev) | Mid_E(KeV) | G (cm^2 sr) | G-small | efficiency
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 0 | p00 | 48.8743 | 56.74 | 52.8072 | 0.0018344 | 0.0018344 | 0.035249
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 1 | p01 | 56.74 | 65.87 | 61.305 | 0.0018344 | 0.0018344 | 0.045544
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 2 | p02 | 65.87 | 76.46 | 71.165 | 0.0018344 | 0.0018344 | 0.078472
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 3 | p03 | 76.46 | 88.77 | 82.615 | 0.0018344 | 0.0018344 | 0.11265
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 4 | p04 | 88.77 | 98.05 | 93.41 | 0.0018344 | 0.0018344 | 0.13768
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 5 | p05 | 98.05 | 108.3 | 103.175 | 0.0018344 | 0.0018344 | 0.14774
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 6 | p06 | 108.3 | 119.63 | 113.965 | 0.0018344 | 0.0018344 | 0.163
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 7 | p07 | 119.63 | 132.14 | 125.885 | 0.0018344 | 0.0018344 | 0.18919
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 8 | p08 | 132.14 | 145.96 | 139.05 | 0.0018344 | 0.0018344 | 0.20554
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 9 | p09 | 145.96 | 161.22 | 153.59 | 0.0018344 | 0.0018344 | 0.20159
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T0_Counts | 10 | p10 | 161.22 | 177 | 169.11 | 0.0018344 | 0.0018344 | 0.19774
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 11 | p11 | 6.6863 | 7.7621 | 7.2242 | 0.0018344 | 0.0018344 | 0.020283
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 12 | p12 | 7.7621 | 8.5738 | 8.1679 | 0.0018344 | 0.0018344 | 0.020209
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 13 | p13 | 8.5738 | 9.4704 | 9.0221 | 0.0018344 | 0.0018344 | 0.021933
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 14 | p14 | 9.4704 | 10.4607 | 9.9656 | 0.0018344 | 0.0018344 | 0.026259
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 15 | p15 | 10.4607 | 11.5547 | 11.0077 | 0.0018344 | 0.0018344 | 0.033807
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 16 | p16 | 11.5547 | 12.763 | 12.1588 | 0.0018344 | 0.0018344 | 0.042612
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 17 | p17 | 12.763 | 14.0976 | 13.4303 | 0.0018344 | 0.0018344 | 0.053096
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 18 | p18 | 14.0976 | 15.5719 | 14.8348 | 0.0018344 | 0.0018344 | 0.067887
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 19 | p19 | 15.5719 | 17.2003 | 16.3861 | 0.0018344 | 0.0018344 | 0.080596
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 20 | p20 | 17.2003 | 18.999 | 18.0997 | 0.0018344 | 0.0018344 | 0.092214
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 21 | p21 | 18.999 | 20.9858 | 19.9924 | 0.0018344 | 0.0018344 | 0.11027
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 22 | p22 | 20.9858 | 23.1804 | 22.0831 | 0.0018344 | 0.0018344 | 0.11788
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 23 | p23 | 23.1804 | 25.6044 | 24.3924 | 0.0018344 | 0.0018344 | 0.1269
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 24 | p24 | 25.6044 | 28.282 | 26.9432 | 0.0018344 | 0.0018344 | 0.13768
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 25 | p25 | 28.282 | 31.2395 | 29.7608 | 0.0018344 | 0.0018344 | 0.15143
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 26 | p26 | 31.2395 | 34.5064 | 32.873 | 0.0018344 | 0.0018344 | 0.20076
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 27 | p27 | 34.5064 | 38.1148 | 36.3106 | 0.0018344 | 0.0018344 | 0.20579
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 28 | p28 | 38.1148 | 42.1006 | 40.1077 | 0.0018344 | 0.0018344 | 0.20399
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 29 | p29 | 42.1006 | 46.5033 | 44.302 | 0.0018344 | 0.0018344 | 0.18304
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T0_Counts | 30 | p30 | 46.5033 | 51.3663 | 48.9348 | 0.0018344 | 0.0018344 | 0.22442
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 414734467.1853 | 662731267.1829 | X | T0_Counts | 31 | p31 | 51.3663 | 56 | 53.6831 | 0.0018344 | 0.0018344 | 0.7
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 0 | p00 | 48.8743 | 56.74 | 52.8072 | 0.002 | 0.002 | 0.035249
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 1 | p01 | 56.74 | 65.87 | 61.305 | 0.002 | 0.002 | 0.045544
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 2 | p02 | 65.87 | 76.46 | 71.165 | 0.002 | 0.002 | 0.078472
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 3 | p03 | 76.46 | 88.77 | 82.615 | 0.002 | 0.002 | 0.11265
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 4 | p04 | 88.77 | 98.05 | 93.41 | 0.002 | 0.002 | 0.13768
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 5 | p05 | 98.05 | 108.3 | 103.175 | 0.002 | 0.002 | 0.14774
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 6 | p06 | 108.3 | 119.63 | 113.965 | 0.002 | 0.002 | 0.163
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 7 | p07 | 119.63 | 132.14 | 125.885 | 0.002 | 0.002 | 0.18919
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 8 | p08 | 132.14 | 145.96 | 139.05 | 0.002 | 0.002 | 0.20554
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 9 | p09 | 145.96 | 161.22 | 153.59 | 0.002 | 0.002 | 0.20159
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T1_Counts | 10 | p10 | 161.22 | 177 | 169.11 | 0.002 | 0.002 | 0.19774
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 11 | p11 | 6.6863 | 7.7621 | 7.2242 | 0.002 | 0.002 | 0.020283
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 12 | p12 | 7.7621 | 8.5738 | 8.1679 | 0.002 | 0.002 | 0.020209
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 13 | p13 | 8.5738 | 9.4704 | 9.0221 | 0.002 | 0.002 | 0.021933
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 14 | p14 | 9.4704 | 10.4607 | 9.9656 | 0.002 | 0.002 | 0.026259
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 15 | p15 | 10.4607 | 11.5547 | 11.0077 | 0.002 | 0.002 | 0.033807
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 16 | p16 | 11.5547 | 12.763 | 12.1588 | 0.002 | 0.002 | 0.042612
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 17 | p17 | 12.763 | 14.0976 | 13.4303 | 0.002 | 0.002 | 0.053096
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 18 | p18 | 14.0976 | 15.5719 | 14.8348 | 0.002 | 0.002 | 0.067887
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 19 | p19 | 15.5719 | 17.2003 | 16.3861 | 0.002 | 0.002 | 0.080596
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 20 | p20 | 17.2003 | 18.999 | 18.0997 | 0.002 | 0.002 | 0.092214
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 21 | p21 | 18.999 | 20.9858 | 19.9924 | 0.002 | 0.002 | 0.11027
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 22 | p22 | 20.9858 | 23.1804 | 22.0831 | 0.002 | 0.002 | 0.11788
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 23 | p23 | 23.1804 | 25.6044 | 24.3924 | 0.002 | 0.002 | 0.1269
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 24 | p24 | 25.6044 | 28.282 | 26.9432 | 0.002 | 0.002 | 0.13768
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 25 | p25 | 28.282 | 31.2395 | 29.7608 | 0.002 | 0.002 | 0.15143
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 26 | p26 | 31.2395 | 34.5064 | 32.873 | 0.002 | 0.002 | 0.20076
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 27 | p27 | 34.5064 | 38.1148 | 36.3106 | 0.002 | 0.002 | 0.20579
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 28 | p28 | 38.1148 | 42.1006 | 40.1077 | 0.002 | 0.002 | 0.20399
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 29 | p29 | 42.1006 | 46.5033 | 44.302 | 0.002 | 0.002 | 0.18304
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T1_Counts | 30 | p30 | 46.5033 | 51.3663 | 48.9348 | 0.002 | 0.002 | 0.22442
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 414734467.1853 | 662731267.1829 | X | T1_Counts | 31 | p31 | 51.3663 | 56 | 53.6831 | 0.002 | 0.002 | 0.7
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 0 | p00 | 48.8743 | 56.74 | 52.8072 | 0.0025083 | 0.0025083 | 0.035249
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 1 | p01 | 56.74 | 65.87 | 61.305 | 0.0025083 | 0.0025083 | 0.045544
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 2 | p02 | 65.87 | 76.46 | 71.165 | 0.0025083 | 0.0025083 | 0.078472
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 3 | p03 | 76.46 | 88.77 | 82.615 | 0.0025083 | 0.0025083 | 0.11265
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 4 | p04 | 88.77 | 98.05 | 93.41 | 0.0025083 | 0.0025083 | 0.13768
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 5 | p05 | 98.05 | 108.3 | 103.175 | 0.0025083 | 0.0025083 | 0.14774
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 6 | p06 | 108.3 | 119.63 | 113.965 | 0.0025083 | 0.0025083 | 0.163
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 7 | p07 | 119.63 | 132.14 | 125.885 | 0.0025083 | 0.0025083 | 0.18919
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 8 | p08 | 132.14 | 145.96 | 139.05 | 0.0025083 | 0.0025083 | 0.20554
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 9 | p09 | 145.96 | 161.22 | 153.59 | 0.0025083 | 0.0025083 | 0.20159
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T2_Counts | 10 | p10 | 161.22 | 177 | 169.11 | 0.0025083 | 0.0025083 | 0.19774
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 11 | p11 | 6.6863 | 7.7621 | 7.2242 | 0.0025083 | 0.0025083 | 0.020283
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 12 | p12 | 7.7621 | 8.5738 | 8.1679 | 0.0025083 | 0.0025083 | 0.020209
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 13 | p13 | 8.5738 | 9.4704 | 9.0221 | 0.0025083 | 0.0025083 | 0.021933
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 14 | p14 | 9.4704 | 10.4607 | 9.9656 | 0.0025083 | 0.0025083 | 0.026259
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 15 | p15 | 10.4607 | 11.5547 | 11.0077 | 0.0025083 | 0.0025083 | 0.033807
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 16 | p16 | 11.5547 | 12.763 | 12.1588 | 0.0025083 | 0.0025083 | 0.042612
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 17 | p17 | 12.763 | 14.0976 | 13.4303 | 0.0025083 | 0.0025083 | 0.053096
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 18 | p18 | 14.0976 | 15.5719 | 14.8348 | 0.0025083 | 0.0025083 | 0.067887
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 19 | p19 | 15.5719 | 17.2003 | 16.3861 | 0.0025083 | 0.0025083 | 0.080596
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 20 | p20 | 17.2003 | 18.999 | 18.0997 | 0.0025083 | 0.0025083 | 0.092214
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 21 | p21 | 18.999 | 20.9858 | 19.9924 | 0.0025083 | 0.0025083 | 0.11027
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 22 | p22 | 20.9858 | 23.1804 | 22.0831 | 0.0025083 | 0.0025083 | 0.11788
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 23 | p23 | 23.1804 | 25.6044 | 24.3924 | 0.0025083 | 0.0025083 | 0.1269
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 24 | p24 | 25.6044 | 28.282 | 26.9432 | 0.0025083 | 0.0025083 | 0.13768
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 25 | p25 | 28.282 | 31.2395 | 29.7608 | 0.0025083 | 0.0025083 | 0.15143
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 26 | p26 | 31.2395 | 34.5064 | 32.873 | 0.0025083 | 0.0025083 | 0.20076
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 27 | p27 | 34.5064 | 38.1148 | 36.3106 | 0.0025083 | 0.0025083 | 0.20579
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 28 | p28 | 38.1148 | 42.1006 | 40.1077 | 0.0025083 | 0.0025083 | 0.20399
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 29 | p29 | 42.1006 | 46.5033 | 44.302 | 0.0025083 | 0.0025083 | 0.18304
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T2_Counts | 30 | p30 | 46.5033 | 51.3663 | 48.9348 | 0.0025083 | 0.0025083 | 0.22442
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 414734467.1853 | 662731267.1829 | X | T2_Counts | 31 | p31 | 51.3663 | 56 | 53.6831 | 0.0025083 | 0.0025083 | 0.7
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 0 | p00 | 48.8743 | 56.74 | 52.8072 | 0.0013622 | 0.0013622 | 0.035249
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 1 | p01 | 56.74 | 65.87 | 61.305 | 0.0013622 | 0.0013622 | 0.045544
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 2 | p02 | 65.87 | 76.46 | 71.165 | 0.0013622 | 0.0013622 | 0.078472
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 3 | p03 | 76.46 | 88.77 | 82.615 | 0.0013622 | 0.0013622 | 0.11265
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 4 | p04 | 88.77 | 98.05 | 93.41 | 0.0013622 | 0.0013622 | 0.13768
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 5 | p05 | 98.05 | 108.3 | 103.175 | 0.0013622 | 0.0013622 | 0.14774
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 6 | p06 | 108.3 | 119.63 | 113.965 | 0.0013622 | 0.0013622 | 0.163
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 7 | p07 | 119.63 | 132.14 | 125.885 | 0.0013622 | 0.0013622 | 0.18919
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 8 | p08 | 132.14 | 145.96 | 139.05 | 0.0013622 | 0.0013622 | 0.20554
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 9 | p09 | 145.96 | 161.22 | 153.59 | 0.0013622 | 0.0013622 | 0.20159
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T3_Counts | 10 | p10 | 161.22 | 177 | 169.11 | 0.0013622 | 0.0013622 | 0.19774
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 11 | p11 | 6.6863 | 7.7621 | 7.2242 | 0.0013622 | 0.0013622 | 0.020283
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 12 | p12 | 7.7621 | 8.5738 | 8.1679 | 0.0013622 | 0.0013622 | 0.020209
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 13 | p13 | 8.5738 | 9.4704 | 9.0221 | 0.0013622 | 0.0013622 | 0.021933
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 14 | p14 | 9.4704 | 10.4607 | 9.9656 | 0.0013622 | 0.0013622 | 0.026259
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 15 | p15 | 10.4607 | 11.5547 | 11.0077 | 0.0013622 | 0.0013622 | 0.033807
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 16 | p16 | 11.5547 | 12.763 | 12.1588 | 0.0013622 | 0.0013622 | 0.042612
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 17 | p17 | 12.763 | 14.0976 | 13.4303 | 0.0013622 | 0.0013622 | 0.053096
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 18 | p18 | 14.0976 | 15.5719 | 14.8348 | 0.0013622 | 0.0013622 | 0.067887
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 19 | p19 | 15.5719 | 17.2003 | 16.3861 | 0.0013622 | 0.0013622 | 0.080596
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 20 | p20 | 17.2003 | 18.999 | 18.0997 | 0.0013622 | 0.0013622 | 0.092214
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 21 | p21 | 18.999 | 20.9858 | 19.9924 | 0.0013622 | 0.0013622 | 0.11027
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 22 | p22 | 20.9858 | 23.1804 | 22.0831 | 0.0013622 | 0.0013622 | 0.11788
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 23 | p23 | 23.1804 | 25.6044 | 24.3924 | 0.0013622 | 0.0013622 | 0.1269
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 24 | p24 | 25.6044 | 28.282 | 26.9432 | 0.0013622 | 0.0013622 | 0.13768
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 25 | p25 | 28.282 | 31.2395 | 29.7608 | 0.0013622 | 0.0013622 | 0.15143
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 26 | p26 | 31.2395 | 34.5064 | 32.873 | 0.0013622 | 0.0013622 | 0.20076
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 27 | p27 | 34.5064 | 38.1148 | 36.3106 | 0.0013622 | 0.0013622 | 0.20579
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 28 | p28 | 38.1148 | 42.1006 | 40.1077 | 0.0013622 | 0.0013622 | 0.20399
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 29 | p29 | 42.1006 | 46.5033 | 44.302 | 0.0013622 | 0.0013622 | 0.18304
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T3_Counts | 30 | p30 | 46.5033 | 51.3663 | 48.9348 | 0.0013622 | 0.0013622 | 0.22442
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 414734467.1853 | 662731267.1829 | X | T3_Counts | 31 | p31 | 51.3663 | 56 | 53.6831 | 0.0013622 | 0.0013622 | 0.7
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 0 | p00 | 48.8743 | 56.74 | 52.8072 | 0.0019716 | 0.0019716 | 0.035249
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 1 | p01 | 56.74 | 65.87 | 61.305 | 0.0019716 | 0.0019716 | 0.045544
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 2 | p02 | 65.87 | 76.46 | 71.165 | 0.0019716 | 0.0019716 | 0.078472
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 3 | p03 | 76.46 | 88.77 | 82.615 | 0.0019716 | 0.0019716 | 0.11265
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 4 | p04 | 88.77 | 98.05 | 93.41 | 0.0019716 | 0.0019716 | 0.13768
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 5 | p05 | 98.05 | 108.3 | 103.175 | 0.0019716 | 0.0019716 | 0.14774
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 6 | p06 | 108.3 | 119.63 | 113.965 | 0.0019716 | 0.0019716 | 0.163
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 7 | p07 | 119.63 | 132.14 | 125.885 | 0.0019716 | 0.0019716 | 0.18919
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 8 | p08 | 132.14 | 145.96 | 139.05 | 0.0019716 | 0.0019716 | 0.20554
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 9 | p09 | 145.96 | 161.22 | 153.59 | 0.0019716 | 0.0019716 | 0.20159
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T4_Counts | 10 | p10 | 161.22 | 177 | 169.11 | 0.0019716 | 0.0019716 | 0.19774
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 11 | p11 | 6.6863 | 7.7621 | 7.2242 | 0.0019716 | 0.0019716 | 0.020283
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 12 | p12 | 7.7621 | 8.5738 | 8.1679 | 0.0019716 | 0.0019716 | 0.020209
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 13 | p13 | 8.5738 | 9.4704 | 9.0221 | 0.0019716 | 0.0019716 | 0.021933
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 14 | p14 | 9.4704 | 10.4607 | 9.9656 | 0.0019716 | 0.0019716 | 0.026259
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 15 | p15 | 10.4607 | 11.5547 | 11.0077 | 0.0019716 | 0.0019716 | 0.033807
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 16 | p16 | 11.5547 | 12.763 | 12.1588 | 0.0019716 | 0.0019716 | 0.042612
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 17 | p17 | 12.763 | 14.0976 | 13.4303 | 0.0019716 | 0.0019716 | 0.053096
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 18 | p18 | 14.0976 | 15.5719 | 14.8348 | 0.0019716 | 0.0019716 | 0.067887
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 19 | p19 | 15.5719 | 17.2003 | 16.3861 | 0.0019716 | 0.0019716 | 0.080596
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 20 | p20 | 17.2003 | 18.999 | 18.0997 | 0.0019716 | 0.0019716 | 0.092214
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 21 | p21 | 18.999 | 20.9858 | 19.9924 | 0.0019716 | 0.0019716 | 0.11027
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 22 | p22 | 20.9858 | 23.1804 | 22.0831 | 0.0019716 | 0.0019716 | 0.11788
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 23 | p23 | 23.1804 | 25.6044 | 24.3924 | 0.0019716 | 0.0019716 | 0.1269
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 24 | p24 | 25.6044 | 28.282 | 26.9432 | 0.0019716 | 0.0019716 | 0.13768
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 25 | p25 | 28.282 | 31.2395 | 29.7608 | 0.0019716 | 0.0019716 | 0.15143
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 26 | p26 | 31.2395 | 34.5064 | 32.873 | 0.0019716 | 0.0019716 | 0.20076
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 27 | p27 | 34.5064 | 38.1148 | 36.3106 | 0.0019716 | 0.0019716 | 0.20579
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 28 | p28 | 38.1148 | 42.1006 | 40.1077 | 0.0019716 | 0.0019716 | 0.20399
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 29 | p29 | 42.1006 | 46.5033 | 44.302 | 0.0019716 | 0.0019716 | 0.18304
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T4_Counts | 30 | p30 | 46.5033 | 51.3663 | 48.9348 | 0.0019716 | 0.0019716 | 0.22442
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 414734467.1853 | 662731267.1829 | X | T4_Counts | 31 | p31 | 51.3663 | 56 | 53.6831 | 0.0019716 | 0.0019716 | 0.7
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 0 | p00 | 48.8743 | 56.74 | 52.8072 | 0.0011646 | 0.0011646 | 0.035249
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 1 | p01 | 56.74 | 65.87 | 61.305 | 0.0011646 | 0.0011646 | 0.045544
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 2 | p02 | 65.87 | 76.46 | 71.165 | 0.0011646 | 0.0011646 | 0.078472
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 3 | p03 | 76.46 | 88.77 | 82.615 | 0.0011646 | 0.0011646 | 0.11265
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 4 | p04 | 88.77 | 98.05 | 93.41 | 0.0011646 | 0.0011646 | 0.13768
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 5 | p05 | 98.05 | 108.3 | 103.175 | 0.0011646 | 0.0011646 | 0.14774
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 6 | p06 | 108.3 | 119.63 | 113.965 | 0.0011646 | 0.0011646 | 0.163
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 7 | p07 | 119.63 | 132.14 | 125.885 | 0.0011646 | 0.0011646 | 0.18919
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 8 | p08 | 132.14 | 145.96 | 139.05 | 0.0011646 | 0.0011646 | 0.20554
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 9 | p09 | 145.96 | 161.22 | 153.59 | 0.0011646 | 0.0011646 | 0.20159
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | O | T5_Counts | 10 | p10 | 161.22 | 177 | 169.11 | 0.0011646 | 0.0011646 | 0.19774
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 11 | p11 | 6.6863 | 7.7621 | 7.2242 | 0.0011646 | 0.0011646 | 0.020283
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 12 | p12 | 7.7621 | 8.5738 | 8.1679 | 0.0011646 | 0.0011646 | 0.020209
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 13 | p13 | 8.5738 | 9.4704 | 9.0221 | 0.0011646 | 0.0011646 | 0.021933
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 14 | p14 | 9.4704 | 10.4607 | 9.9656 | 0.0011646 | 0.0011646 | 0.026259
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 15 | p15 | 10.4607 | 11.5547 | 11.0077 | 0.0011646 | 0.0011646 | 0.033807
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 16 | p16 | 11.5547 | 12.763 | 12.1588 | 0.0011646 | 0.0011646 | 0.042612
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 17 | p17 | 12.763 | 14.0976 | 13.4303 | 0.0011646 | 0.0011646 | 0.053096
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 18 | p18 | 14.0976 | 15.5719 | 14.8348 | 0.0011646 | 0.0011646 | 0.067887
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 19 | p19 | 15.5719 | 17.2003 | 16.3861 | 0.0011646 | 0.0011646 | 0.080596
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 20 | p20 | 17.2003 | 18.999 | 18.0997 | 0.0011646 | 0.0011646 | 0.092214
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 21 | p21 | 18.999 | 20.9858 | 19.9924 | 0.0011646 | 0.0011646 | 0.11027
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 22 | p22 | 20.9858 | 23.1804 | 22.0831 | 0.0011646 | 0.0011646 | 0.11788
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 23 | p23 | 23.1804 | 25.6044 | 24.3924 | 0.0011646 | 0.0011646 | 0.1269
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 24 | p24 | 25.6044 | 28.282 | 26.9432 | 0.0011646 | 0.0011646 | 0.13768
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 25 | p25 | 28.282 | 31.2395 | 29.7608 | 0.0011646 | 0.0011646 | 0.15143
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 26 | p26 | 31.2395 | 34.5064 | 32.873 | 0.0011646 | 0.0011646 | 0.20076
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 27 | p27 | 34.5064 | 38.1148 | 36.3106 | 0.0011646 | 0.0011646 | 0.20579
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 28 | p28 | 38.1148 | 42.1006 | 40.1077 | 0.0011646 | 0.0011646 | 0.20399
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 29 | p29 | 42.1006 | 46.5033 | 44.302 | 0.0011646 | 0.0011646 | 0.18304
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 347112066.1839 | 662731267.1829 | H | T5_Counts | 30 | p30 | 46.5033 | 51.3663 | 48.9348 | 0.0011646 | 0.0011646 | 0.22442
TOFxPH Hydrogen - High Energy Res Low Time Res | TOFxPH_H_HELT | 414734467.1853 | 662731267.1829 | X | T5_Counts | 31 | p31 | 51.3663 | 56 | 53.6831 | 0.0011646 | 0.0011646 | 0.7

## SW
Product Description | Product Name | StartET | StopET | Species | Field Name | Channel | ChName | E_Low (Kev) | E_Hi (Kev) | Mid_E(KeV) | G (cm^2 sr) | G-small | efficiency
Space Weather Rates | SW | 414734467.1853 | 662731267.1829 | Ions | Fast0_Counts | 0 | swfa00 | 49 | 60 | 55 | 0.0013716 | 6.858e-05 | 1
Space Weather Rates | SW | 347112066.1839 | 662731267.1829 | Ions | Fast1_Counts | 0 | swfb00 | 89.9 | 109.7 | 100 | 0.0013716 | 6.858e-05 | 1
Space Weather Rates | SW | 347112066.1839 | 662731267.1829 | Ions | Fast2_Counts | 0 | swfc00 | 133.7 | 163 | 148 | 0.0013716 | 6.858e-05 | 1
Space Weather Rates | SW | 347112066.1839 | 662731267.1829 | Ions | Fast3_Counts | 0 | swfd00 | 242.4 | 295.6 | 269 | 0.0013716 | 6.858e-05 | 1
Space Weather Rates | SW | 347112066.1839 | 662731267.1829 | Ions | Slow0_Counts | 0 | swsa00 | 972.4925 | 1071.6616 | 1021 | 0.0013716 | 6.858e-05 | 1
Space Weather Rates | SW | 347112066.1839 | 662731267.1829 | He | Slow1_Counts | 0 | swsb00 | 9074.623 | 10000 | 9537.3115 | 0.0013716 | 6.858e-05 | 1